## MEXReport
Primary Standards Dosimetry Laboratory, Medical Radiation Services
619 Lower Plenty Road, Yallambie, Victoria 3085, Australia
Tel: +613 9433 2211 Fax: +613 9432 1835
E-mail: psdl@arpansa.gov.au Web: www.arpansa.gov.au
In reply please quote:
None
CAL NUMBER
CALIBRATION REPORT
on a therapy ionisation chamber for
MEDIUM-ENERGY KILOVOLTAGE X-RAYS
Client |  |  |  | Client Name
Address 1
Address 2
Ionisation chamber |  |  |  | Chamber Info
Period of tests |  |  |  | Date of measurement
Previous calibration |  |  |  | Not previously calibrated at ARPANSA | [11 empty] | IRC-FORM-2100J (ARPANSA-FORM-1856)
Test By |  |  |  | Test By
Report By |  |  |  | Report By
Report date |  |  |  | Report Date
Direct inquiries to |  |  |  | Chris Oliver
Tel: (03) 9433 2333 |  |  |  | Email: psdl@arpansa.gov.au
Signed: |  |  |  |  | (Authorised Signatory) | [5 empty] | Date:
Chris Oliver, Senior Radiation Scientist, Primary Standards Dosimetry Laboratory
per C-M Larsson, CEO of ARPANSA
This certificate is consistent with the capabilities that are
included in Appendix C of the MRA drawn up by the
|  |  | Accredited for compliance with ISO/IEC 17025 - Calibration. | [7 empty] | CIPM. Under the MRA, all participating institutes
|  |  | The results of the tests, calibrations and/or | [7 empty] | recognize the validity of each other’s calibration and
|  |  | measurements included in this document are | [7 empty] | measurement certificates for the quantities, ranges and
|  |  | traceable to Australian/national standards | [7 empty] | measurement uncertainties specified in Appendix C (for
|  |  | NATA Accredited Laboratory Number: 14433 | [7 empty] | details see www.bipm.org).
Calibration No: + CAL NUMBER | [5 empty] | Calibration No: + CAL NUMBER |  |  |  |  | Checked: |  |  |  |  | Page 1 of 6
GENERAL COMMENTS
Preliminary Inspection
– | The ionisation chamber had no obvious damage or faults on receipt.
Calibration Coefficient
– | The calibration coefficient is the number by which the charge from the chamber, in nC, must be multiplied to obtain the air kerma [1]. The calibration factor for the electrometer must also be taken into account when measuring the charge from the chamber.
Calibration Coefficients for Medium-Energy X-ray (MEX) Qualities in Air
– | The calibration coefficients for the chamber for each X-ray beam quality from the Gulmay Comet X-ray generator were determined by comparison with the ARPANSA Medium Energy Free-Air Chamber, which is the Australian primary standard of air kerma for medium energy X-rays.
– | The Gulmay Comet X-ray generator is constant potential and the X-ray tube has a tungsten target.
Recombination Correction Measurement
– | Not measured.
Polarity Correction Measurement
– | Not measured.
Notes
– | The ionisation chamber was tested in accordance with ARPANSA Standard Operational Procedure ARPANSA-SOP-0816 Version 7.
References
[1] | AAPM protocol for 40–300 kV x-ray beam dosimetry in radiotherapy and radiobiology, C.-M. Ma, Chair, C. W. Coffey, L. A. DeWerd, C. Liu, R. Nath, S. M. Seltzer, J. P. Seuntjens, Med. Phys. 28 (6) 868-893, 2001
Calibration No: + CAL NUMBER | [5 empty] | Calibration No: + CAL NUMBER |  |  |  |  | Checked: |  |  |  |  | Page 2 of 6
Air Kerma Calibration Certificate - Medium-Energy X-rays
Client |  |  |  | Report Date
Ionisation chamber |  |  |  | Chamber Info
Polarising voltage |  |  |  | Variable + on the guard electrode
Collected charge polarity |  |  |  | Positive + (Central Electrode Negative
Reference point |  |  |  | The geometrical centre of the cavity
Geometry |  |  |  | Mark on chamber stem facing the radiation source
Chamber stem vertically upwards, cable down
Horizontal radiation beam
Source-detector distance 100 cm
Circular beam of diameter 10 cm
Build-up |  |  |  | Build-up cap removed except where stated. Calibrated free in air.
Polarity and recombination |  |  |  | Corrections not applied
Reference conditions |  |  |  | 20oC, 101.325 kPa and 50% humidity
Measurement date(s) |  |  |  | Date of measurement
Uncertainties (U) are given at a confidence level of approximately 95% (k=2)
Table 1: Subset of air kerma calibration coefficients
Beam code |  | Tube voltage | Added filter | Added filter |  | HVL |  | HVL |  | Nominal effective energy [1] |  | Nominal air kerma rate |  | NK [2] | U
|  | kV | mm Al | mm Cu |  | mm Al |  | mm Cu |  | keV |  | mGy/s |  | mGy/nC | %
NXA50 |  | 50 | 4 |  |  | 2.3884 |  | 0.07999 |  | 29.934 |  | 0.36147 |  | 45.0217 | 1.422
NXA70 |  | 70 | 4 |  |  | 3.193 |  | 0.11136 |  | 34.275 |  | 0.80409 |  | 44.9749 | 1.4211
NXB100 |  | 100 | 4.5 |  |  | 4.7352 |  | 0.18269 |  | 41.587 |  | 1.5384 |  | 44.7785 | 1.404
NXC120 |  | 120 | 6 |  |  | 6.377 |  | 0.28062 |  | 48.945 |  | 1.7719 |  | 44.5963 | 1.4061
NXD140 |  | 140 | 9 |  |  | 8.436 |  | 0.45082 |  | 58.063 |  | 1.758 |  | 44.385 | 1.4143
NXE150 |  | 150 | 4 | 0.5 |  |  |  | 0.83812 |  | 72.15 |  | 1.4008 |  | 44.0399 | 1.4253
NXF200 |  | 200 | 4 | 1 |  |  |  | 1.6297 |  | 94.739 |  | 1.5624 |  | 43.6904 | 1.4009
NXG250 |  | 250 | 4 | 1.6 |  |  |  | 2.5712 |  | 120.097 |  | 1.7673 |  | 43.5153 | 1.4182
NXH280 |  | 280 | 4 | 3 |  |  |  | 3.5006 |  | 147.05 |  | 1.3697 |  | 43.3321 | 1.4392
NXH300 |  | 300 | 4 | 3 |  |  |  | 3.7048 |  | 153.21 |  | 1.6753 |  | 43.2701 | 1.5467
NXH300* |  | 300 | 4 | 3 |  |  |  | 3.7048 |  | 153.21 |  | 1.6756 |  | 43.2738 | 1.4099
[1] The energy of a monoenergetic beam with the same HVL in mm of Cu
[2] The air kerma calibration coefficient
* With buildup cap on
Calibrated by: Operator
Calibration No: + CAL NUMBER | [5 empty] | Calibration No: + CAL NUMBER |  |  |  |  | Checked: |  |  |  |  | Page 3 of 6
Table 2: Complete set of air kerma calibration coefficients for all MEX beams
Beam code |  | Tube voltage | Added filter | Added filter |  | HVL |  | HVL |  | Nominal effective energy [1] |  | Nominal air kerma rate |  | NK [2] | U
|  | kV | mm Al | mm Cu |  | mm Al |  | mm Cu |  | keV |  | mGy/s |  | mGy/nC | %
NXJ40 |  | 40 | 0.5 |  |  | 0.57493 |  |  |  | 17.956 |  | 1.9269 |  | 47.6935 | 1.5315
NXK40 |  | 40 | 1 |  |  | 0.92589 |  |  |  | 20.93 |  | 1.0342 |  | 46.2033 | 1.4416
NXA40 |  | 40 | 4 |  |  | 1.8979 |  | 0.062723 |  | 27.068 |  | 0.18221 |  | 44.9989 | 1.4338
NXJ50 |  | 50 | 0.5 |  |  | 0.6841 |  |  |  | 19.5 |  | 2.6953 |  | 47.3162 | 1.4093
NXK50 |  | 50 | 1 |  |  | 1.1174 |  |  |  | 22.646 |  | 1.5613 |  | 45.9148 | 1.4897
NXA50 |  | 50 | 4 |  |  | 2.3884 |  | 0.07999 |  | 29.934 |  | 0.36147 |  | 45.0217 | 1.422
NXB50 |  | 50 | 4.5 |  |  | 2.5293 |  | 0.084996 |  | 30.623 |  | 0.30801 |  | 44.9092 | 1.5131
NXJ60 |  | 60 | 0.5 |  |  | 0.78881 |  |  |  | 20.857 |  | 3.4163 |  | 46.8133 | 1.5437
NXK60 |  | 60 | 1 |  |  | 1.2886 |  |  |  | 24.095 |  | 2.0865 |  | 45.7921 | 1.4909
NXA60 |  | 60 | 4 |  |  | 2.8059 |  | 0.0958 |  | 32.251 |  | 0.57234 |  | 45.048 | 1.4116
NXJ70 |  | 70 | 0.5 |  |  | 0.89521 |  |  |  | 22.147 |  | 4.0878 |  | 46.5843 | 1.4198
NXK70 |  | 70 | 1 |  |  | 1.4584 |  |  |  | 25.442 |  | 2.6003 |  | 45.6259 | 1.4131
NXA70 |  | 70 | 4 |  |  | 3.193 |  | 0.11136 |  | 34.275 |  | 0.80409 |  | 44.9749 | 1.4211
NXB70 |  | 70 | 4.5 |  |  | 3.3916 |  | 0.11943 |  | 35.19 |  | 0.709 |  | 44.9306 | 1.4134
NXC70 |  | 70 | 6 |  |  | 3.9521 |  | 0.14141 |  | 37.509 |  | 0.50734 |  | 44.8346 | 1.5176
NXJ80 |  | 80 | 0.5 |  |  | 1.0169 |  |  |  | 23.48 |  | 4.7438 |  | 46.3561 | 1.4502
NXK80 |  | 80 | 1 |  |  | 1.6432 |  |  |  | 26.842 |  | 3.1296 |  | 45.5169 | 1.4293
NXA80 |  | 80 | 4 |  |  | 3.6197 |  | 0.12924 |  | 36.328 |  | 1.072 |  | 44.8842 | 1.405
NXJ90 |  | 90 | 0.5 |  |  | 1.1525 |  |  |  | 24.878 |  | 5.3833 |  | 46.158 | 1.4322
NXK90 |  | 90 | 1 |  |  | 1.8393 |  |  |  | 28.32 |  | 3.6673 |  | 45.3793 | 1.4209
NXA90 |  | 90 | 4 |  |  | 4.0394 |  | 0.14848 |  | 38.385 |  | 1.3709 |  | 44.8408 | 1.4118
NXJ100 |  | 100 | 0.5 |  |  | 1.3031 |  |  |  | 26.34 |  | 6.0072 |  | 45.9748 | 1.4127
NXK100 |  | 100 | 1 |  |  | 2.0481 |  |  |  | 29.867 |  | 4.2111 |  | 45.3403 | 1.4193
NXB100 |  | 100 | 4.5 |  |  | 4.7352 |  | 0.18269 |  | 41.587 |  | 1.5384 |  | 44.7785 | 1.404
NXC100 |  | 100 | 6 |  |  | 5.485 |  | 0.2214 |  | 44.583 |  | 1.193 |  | 44.7312 | 1.4425
NXD100 |  | 100 | 9 |  |  | 6.6136 |  | 0.29052 |  | 49.084 |  | 0.79055 |  | 44.6326 | 1.4404
NXB120 |  | 120 | 4.5 |  |  | 5.556 |  | 0.22972 |  | 45.744 |  | 2.2136 |  | 44.64 | 1.4134
NXC120 |  | 120 | 6 |  |  | 6.377 |  | 0.28062 |  | 48.945 |  | 1.7719 |  | 44.5963 | 1.4061
NXD120 |  | 120 | 9 |  |  | 7.5882 |  | 0.3681 |  | 53.752 |  | 1.2349 |  | 44.5071 | 1.4107
NXE120 |  | 120 | 4 | 0.5 |  | 10.3105 |  | 0.63484 |  | 65 |  | 0.73864 |  | 44.2405 | 1.4415
NXB140 |  | 140 | 4.5 |  |  | 6.333 |  | 0.28214 |  | 49.868 |  | 2.9683 |  | 44.4724 | 1.4507
NXC140 |  | 140 | 6 |  |  | 7.1952 |  | 0.34468 |  | 53.139 |  | 2.4337 |  | 44.4387 | 1.4083
NXD140 |  | 140 | 9 |  |  | 8.436 |  | 0.45082 |  | 58.063 |  | 1.758 |  | 44.385 | 1.4143
NXE140 |  | 140 | 4 | 0.5 |  |  |  | 0.76915 |  | 69.85 |  | 1.157 |  | 44.1876 | 1.4486
NXF140 |  | 140 | 4 | 1 |  |  |  | 1.0269 |  | 78.519 |  | 0.60047 |  | 43.9116 | 1.4213
NXC150 |  | 150 | 6 |  |  | 7.5822 |  | 0.37861 |  | 55.187 |  | 2.7968 |  | 44.3997 | 1.404
NXD150 |  | 150 | 9 |  |  | 8.8317 |  | 0.49292 |  | 60.128 |  | 2.0497 |  | 44.3357 | 1.4182
NXE150 |  | 150 | 4 | 0.5 |  |  |  | 0.83812 |  | 72.15 |  | 1.4008 |  | 44.0399 | 1.4253
NXF150 |  | 150 | 4 | 1 |  |  |  | 1.1251 |  | 81.358 |  | 0.89307 |  | 43.9346 | 1.5487
NXG150 |  | 150 | 4 | 1.6 |  |  |  | 1.3834 |  | 89.066 |  | 0.44635 |  | 43.692 | 1.4055
…continued p5
Calibration No: + CAL NUMBER | [5 empty] | Calibration No: + CAL NUMBER |  |  |  |  | Checked: |  |  |  |  | Page 4 of 6
Beam code |  | Tube voltage | Added filter | Added filter |  | HVL |  | HVL |  | Nominal effective energy [1] |  | Nominal air kerma rate |  | NK [2] | U
|  | kV | mm Al | mm Cu |  | mm Al |  | mm Cu |  | keV |  | mGy/s |  | mGy/nC | %
NXD200 |  | 200 | 9 |  |  | 10.5331 |  | 0.7323 |  | 69.761 |  | 2.8398 |  | 44.0646 | 1.4218
NXE200 |  | 200 | 4 | 0.5 |  |  |  | 1.2072 |  | 82.81 |  | 2.1995 |  | 43.7836 | 1.4026
NXF200 |  | 200 | 4 | 1 |  |  |  | 1.6297 |  | 94.739 |  | 1.5624 |  | 43.6904 | 1.4009
NXG200 |  | 200 | 4 | 1.6 |  |  |  | 2.0015 |  | 105.252 |  | 1.1613 |  | 43.5761 | 1.4075
NXH200 |  | 200 | 4 | 3 |  |  |  | 2.5415 |  | 120.147 |  | 0.64561 |  | 43.4856 | 1.6172
NXE250 |  | 250 | 4 | 0.5 |  |  |  | 1.6128 |  | 92.84 |  | 2.9839 |  | 43.7298 | 1.422
NXF250 |  | 250 | 4 | 1 |  |  |  | 2.1357 |  | 107.324 |  | 2.2566 |  | 43.5481 | 1.5783
NXG250 |  | 250 | 4 | 1.6 |  |  |  | 2.5712 |  | 120.097 |  | 1.7673 |  | 43.5153 | 1.4182
NXH250 |  | 250 | 4 | 3 |  |  |  | 3.1623 |  | 137.455 |  | 1.1621 |  | 43.3941 | 1.4255
NXI250 |  | 250 | 4 | 5 |  |  |  | 3.5952 |  | 149.507 |  | 0.73254 |  | 43.2622 | 1.4978
NXF280 |  | 280 | 4 | 1 |  |  |  | 2.4332 |  | 114.612 |  | 2.5003 |  | 43.5479 | 1.4412
NXG280 |  | 280 | 4 | 1.6 |  |  |  | 2.8809 |  | 128.495 |  | 2.0035 |  | 43.417 | 1.4112
NXH280 |  | 280 | 4 | 3 |  |  |  | 3.5006 |  | 147.05 |  | 1.3697 |  | 43.3321 | 1.4392
NXI280 |  | 280 | 4 | 5 |  |  |  | 3.9335 |  | 159.918 |  | 0.89372 |  | 43.2496 | 1.4139
NXG300 |  | 300 | 4 | 1.6 |  |  |  | 3.0952 |  | 133.916 |  | 2.3987 |  | 43.3593 | 1.4146
NXH300 |  | 300 | 4 | 3 |  |  |  | 3.7048 |  | 153.21 |  | 1.6753 |  | 43.2701 | 1.5467
NXI300 |  | 300 | 4 | 5 |  |  |  | 4.1454 |  | 166.581 |  | 1.1162 |  | 43.134 | 1.4145
NXH320 |  | 320 | 4 | 3 |  |  |  | 3.8973 |  | 159.217 |  | 1.6173 |  | 43.1318 | 1.442
NXI320 |  | 320 | 4 | 5 |  |  |  | 4.3355 |  | 173.059 |  | 1.0975 |  | 43.0485 | 1.4067
NXG250* |  | 250 | 4 | 1.6 |  |  |  | 2.5712 |  | 120.097 |  | 1.7715 |  | 43.3443 | 1.4048
NXH280* |  | 280 | 4 | 3 |  |  |  | 3.5006 |  | 147.05 |  | 1.3702 |  | 43.3328 | 1.4124
NXH300* |  | 300 | 4 | 3 |  |  |  | 3.7048 |  | 153.21 |  | 1.6756 |  | 43.2738 | 1.4099
[1] The energy of a monoenergetic beam with the same HVL in mm of Cu
[2] The air kerma calibration coefficient
* With buildup cap on
#REF!
Calibration No: + CAL NUMBER | [5 empty] | Calibration No: + CAL NUMBER |  |  |  |  | Checked: |  |  |  |  | Page 5 of 6
#REF!
#REF!
Calibration No: + CAL NUMBER | [5 empty] | Calibration No: + CAL NUMBER |  |  |  |  | Checked: |  |  |  |  | Page 6 of 6