## client
[COMET X-RAY MEASUREMENT]
Filename |  | C:\CRData\12Feb2021-4-31P.csv
Date |  | 12/02/2021
Chamber |  | PTW 30013 5122
Description |  | Standard ratio measurement
Software |  | S:\Medical_Rad\Radiotherapy\Ionizing Radiation\IRS Inhouse Software\Twin Weblines LEX and MEX\Twin Webline Comet v8_0-PRODUCTION.vi
Backgrounds |  | 90
Measurements |  | 30
Trolley (mm) |  | 1090
SCD (mm) |  | 1000
Aperture wheel |  | 2 CM
Comment |  | Set 1 on client chamber
Monitor electrometer range |  | Medium
Monitor HV |  | 300
MEFAC-IC electrometer range |  | Low
IC HV |  | -250
Client name |  | ClientA_Name | [6 empty] | Client identifier string |  |  | *this line may not be present in all files
Address 1 |  | 200 Street Name | [6 empty] | Address line 1 |  |  | *this line may not be present in all files
Address 2 |  | Suburb NSW 2020 | [6 empty] | Address line 2 |  |  | *this line may not be present in all files
Operator |  | Duncan Butler | [6 empty] | Who did the measurements |  |  | *this line may not be present in all files
CAL Number |  | CAL00001 | [6 empty] | ARPANSA Job ID |  |  | *this line may not be present in all files
[DATA]
kV | mA | BarCode | XraysOn | HVLFilter(mm) | Filter | FilterReady | HVLReady | N | Current1(pA) | Current2(pA) | P(kPa) | T(MC) | T(Air) | T(SC) | H(%) | Comment
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 1 | 0.2 | 0.001 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 2 | 0.25 | 0.001 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 3 | 0.15 | 0.001 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 4 | 0.05 | 0.002 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 5 | 0.2 | 0.001 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 6 | 0.15 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 7 | 0.15 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 8 | 0.1 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 9 | 0.2 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 10 | 0.2 | 0.001 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 11 | 0.3 | 0.001 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 12 | 0.15 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 13 | 0.15 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 14 | 0 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 15 | 0.05 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 16 | 0 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 17 | 0.15 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 18 | 0 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 19 | 0 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 20 | 0 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 21 | 0.2 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 22 | 0 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 23 | 0.05 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 24 | 0.15 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 25 | 0.1 | 0 | 100.042 | 22.663052 | 22.573487 | 23.205638 | 58.047587 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 26 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 27 | 0.05 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 28 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 29 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 30 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 31 | 0.05 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 32 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 33 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 34 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 35 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 36 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 37 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 38 | 0.4 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 39 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 40 | 0.35 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 41 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 42 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 43 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 44 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 45 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 46 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 47 | 0.05 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 48 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 49 | 0.1 | 0.001 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 50 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 51 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 52 | 0 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 53 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 54 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 55 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 56 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 57 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 58 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 59 | 0.3 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 60 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 61 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 62 | 0 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 63 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 64 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 65 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 66 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 67 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 68 | 0.05 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 69 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 70 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 71 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 72 | 0 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 73 | 0.05 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 74 | 0.25 | 0.001 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 75 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 76 | 0.35 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 77 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 78 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 79 | 0.25 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 80 | 0 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 81 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 82 | 0 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 83 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 84 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 85 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 86 | 0.1 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 87 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 88 | 0.2 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 89 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 90 | 0.15 | 0 | 100.04 | 22.660328 | 22.564224 | 23.207873 | 58.058631 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 1 | 3713 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 2 | 3707 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 3 | 3710 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 4 | 3712 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 5 | 3709 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 6 | 3714 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 7 | 3708 | -37.04 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 8 | 3710 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 9 | 3709 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 10 | 3711 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 11 | 3710 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 12 | 3708 | -37.04 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 13 | 3713 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 14 | 3709 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 15 | 3713 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 16 | 3708 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 17 | 3709 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 18 | 3715 | -37.05 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 19 | 3708 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 20 | 3716 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 21 | 3710 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 22 | 3712 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 23 | 3712 | -37.06 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 24 | 3711 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 25 | 3713 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 26 | 3708 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 27 | 3712 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 28 | 3709 | -37.07 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 29 | 3711 | -37.08 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 30 | 3711 | -37.08 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 1 | 5373 | -52.59 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 2 | 5378 | -52.61 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 3 | 5373 | -52.61 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 4 | 5374 | -52.62 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 5 | 5374 | -52.62 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 6 | 5375 | -52.61 | 100.042 | 22.654283 | 22.551198 | 23.211482 | 58.067423 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 7 | 5376 | -52.62 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 8 | 5372 | -52.62 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 9 | 5378 | -52.62 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 10 | 5373 | -52.62 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 11 | 5377 | -52.62 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 12 | 5379 | -52.63 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 13 | 5374 | -52.63 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 14 | 5379 | -52.63 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 15 | 5373 | -52.63 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 16 | 5375 | -52.65 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 17 | 5375 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 18 | 5376 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 19 | 5376 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 20 | 5375 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 21 | 5378 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 22 | 5374 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 23 | 5377 | -52.65 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 24 | 5373 | -52.63 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 25 | 5375 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 26 | 5380 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 27 | 5375 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 28 | 5381 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 29 | 5374 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 30 | 5376 | -52.64 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 1 | 6996 | -67.19 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 2 | 6996 | -67.24 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 3 | 6999 | -67.39 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 4 | 6995 | -67.2 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 5 | 6999 | -67.38 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 6 | 6995 | -67.38 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 7 | 6997 | -67.63 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 8 | 7000 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 9 | 6995 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 10 | 7002 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 11 | 6996 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 12 | 6996 | -67.41 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 13 | 6997 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 14 | 6997 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 15 | 6999 | -67.22 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 16 | 6998 | -67.59 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 17 | 6999 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 18 | 6996 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 19 | 6998 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 20 | 6996 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 21 | 6998 | -67.27 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 22 | 7002 | -67.13 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 23 | 6997 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 24 | 7002 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 25 | 6996 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 26 | 6998 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 27 | 6999 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 28 | 6997 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 29 | 7000 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 30 | 6996 | -67.4 | 100.042 | 22.652232 | 22.543511 | 23.21338 | 58.067995 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 1 | 8534 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 2 | 8530 | -81.21 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 3 | 8532 | -81.2 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 4 | 8536 | -81.2 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 5 | 8534 | -81.19 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 6 | 8538 | -81.2 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 7 | 8532 | -81.21 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 8 | 8532 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 9 | 8533 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 10 | 8534 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 11 | 8532 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 12 | 8534 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 13 | 8536 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 14 | 8532 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 15 | 8534 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 16 | 8531 | -81.22 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 17 | 8531 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 18 | 8537 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 19 | 8532 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 20 | 8538 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 21 | 8533 | -81.24 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 22 | 8534 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 23 | 8535 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 24 | 8534 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 25 | 8536 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 26 | 8532 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 27 | 8534 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 28 | 8531 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 29 | 8535 | -81.24 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 30 | 8533 | -81.23 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 1 | 10060 | -94.85 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 2 | 10060 | -94.83 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 3 | 10060 | -94.83 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 4 | 10060 | -94.84 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 5 | 10060 | -94.84 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 6 | 10060 | -94.85 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 7 | 10060 | -95.21 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 8 | 10060 | -94.85 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 9 | 10060 | -94.86 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 10 | 10060 | -94.85 | 100.045 | 22.648418 | 22.532374 | 23.216406 | 58.064473 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 11 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 12 | 10060 | -94.86 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 13 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 14 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 15 | 10060 | -94.84 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 16 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 17 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 18 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 19 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 20 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 21 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 22 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 23 | 10060 | -94.84 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 24 | 10060 | -94.84 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 25 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 26 | 10070 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 27 | 10060 | -94.84 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 28 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 29 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 30 | 10060 | -94.85 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 1 | 11560 | -108.1 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 2 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 3 | 11560 | -108.1 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 4 | 11560 | -108.1 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 5 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 6 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 7 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 8 | 11570 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 9 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 10 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 11 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 12 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 13 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 14 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 15 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 16 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 17 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 18 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 19 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 20 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 21 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 22 | 11570 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 23 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 24 | 11570 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 25 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 26 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 27 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 28 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 29 | 11570 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 30 | 11560 | -108.2 | 100.044 | 22.645015 | 22.525989 | 23.218846 | 58.074522 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 1 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 2 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 3 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 4 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 5 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 6 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 7 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 8 | 13040 | -120.9 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 9 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 10 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 11 | 13040 | -120.9 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 12 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 13 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 14 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 15 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 16 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 17 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 18 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 19 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 20 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 21 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 22 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 23 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 24 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 25 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 26 | 13030 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 27 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 28 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 29 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 30 | 13040 | -121.2 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 1 | 2000 | -20.81 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 2 | 2004 | -20.8 | 100.044 | 22.641389 | 22.515427 | 23.221023 | 58.089596 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 3 | 2001 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 4 | 2001 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 5 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 6 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 7 | 2003 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 8 | 1999 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 9 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 10 | 2000 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 11 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 12 | 2005 | -20.79 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 13 | 2001 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 14 | 2005 | -20.79 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 15 | 2000 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 16 | 2003 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 17 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 18 | 2001 | -20.81 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 19 | 2003 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 20 | 2001 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 21 | 2003 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 22 | 2000 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 23 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 24 | 2001 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 25 | 2002 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 26 | 2005 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 27 | 2000 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 28 | 2004 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 29 | 2000 | -20.8 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 30 | 2003 | -20.81 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 1 | 3178 | -31.68 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 2 | 3177 | -31.68 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 3 | 3180 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 4 | 3176 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 5 | 3180 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 6 | 3177 | -31.7 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 7 | 3178 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 8 | 3181 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 9 | 3177 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 10 | 3181 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 11 | 3176 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 12 | 3177 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 13 | 3178 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 14 | 3178 | -31.68 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 15 | 3179 | -31.68 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 16 | 3178 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 17 | 3179 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 18 | 3177 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 19 | 3177 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 20 | 3178 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 21 | 3178 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 22 | 3181 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 23 | 3177 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 24 | 3181 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 25 | 3177 | -31.69 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 26 | 3179 | -31.7 | 100.044 | 22.641168 | 22.50803 | 23.223021 | 58.074355 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 27 | 3179 | -31.69 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 28 | 3177 | -31.69 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 29 | 3181 | -31.69 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 30 | 3177 | -31.7 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 1 | 4400 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 2 | 4397 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 3 | 4398 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 4 | 4401 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 5 | 4397 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 6 | 4401 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 7 | 4398 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 8 | 4398 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 9 | 4398 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 10 | 4398 | -42.58 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 11 | 4399 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 12 | 4396 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 13 | 4399 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 14 | 4396 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 15 | 4398 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 16 | 4398 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 17 | 4397 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 18 | 4401 | -42.59 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 19 | 4397 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 20 | 4401 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 21 | 4399 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 22 | 4398 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 23 | 4400 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 24 | 4399 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 25 | 4402 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 26 | 4397 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 27 | 4398 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 28 | 4398 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 29 | 4398 | -42.61 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 30 | 4397 | -42.6 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 1 | 5608 | -53.34 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 2 | 5612 | -53.32 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 3 | 5608 | -53.32 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 4 | 5610 | -53.32 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 5 | 5611 | -53.31 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 6 | 5611 | -53.31 | 100.045 | 22.64038 | 22.502342 | 23.224527 | 58.073801 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 7 | 5611 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 8 | 5609 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 9 | 5610 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 10 | 5608 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 11 | 5610 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 12 | 5612 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 13 | 5609 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 14 | 5614 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 15 | 5608 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 16 | 5612 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 17 | 5609 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 18 | 5611 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 19 | 5611 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 20 | 5610 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 21 | 5613 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 22 | 5609 | -53.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 23 | 5611 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 24 | 5609 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 25 | 5610 | -53.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 26 | 5613 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 27 | 5611 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 28 | 5612 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 29 | 5609 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 30 | 5612 | -53.31 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 1 | 6867 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 2 | 6868 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 3 | 6870 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 4 | 6866 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 5 | 6870 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 6 | 6865 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 7 | 6869 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 8 | 6871 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 9 | 6866 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 10 | 6870 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 11 | 6866 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 12 | 6867 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 13 | 6868 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 14 | 6868 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 15 | 6871 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 16 | 6867 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 17 | 6870 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 18 | 6867 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 19 | 6869 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 20 | 6867 | -64.43 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 21 | 6868 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 22 | 6871 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 23 | 6868 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 24 | 6870 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 25 | 6867 | -64.29 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 26 | 6867 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 27 | 6868 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 28 | 6868 | -64.16 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 29 | 6870 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 30 | 6867 | -64.3 | 100.039 | 22.639333 | 22.499409 | 23.227047 | 58.070487 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 1 | 8153 | -75.48 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 2 | 8149 | -75.47 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 3 | 8152 | -75.48 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 4 | 8154 | -75.71 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 5 | 8150 | -75.49 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 6 | 8154 | -75.5 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 7 | 8152 | -75.51 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 8 | 8152 | -75.5 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 9 | 8152 | -75.5 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 10 | 8151 | -75.51 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 11 | 8153 | -75.74 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 12 | 8152 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 13 | 8154 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 14 | 8151 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 15 | 8153 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 16 | 8150 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 17 | 8150 | -75.75 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 18 | 8155 | -75.53 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 19 | 8151 | -75.53 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 20 | 8155 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 21 | 8151 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 22 | 8152 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 23 | 8152 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 24 | 8152 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 25 | 8154 | -75.53 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 26 | 8151 | -75.53 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 27 | 8154 | -75.53 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 28 | 8152 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 29 | 8154 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 30 | 8152 | -75.52 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 1 | 9449 | -86.83 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 2 | 9451 | -86.84 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 3 | 9450 | -86.83 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 4 | 9452 | -86.82 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 5 | 9450 | -86.82 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 6 | 9450 | -86.82 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 7 | 9449 | -86.82 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 8 | 9450 | -86.83 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 9 | 9452 | -86.82 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 10 | 9448 | -86.82 | 100.038 | 22.639584 | 22.497611 | 23.228486 | 58.061492 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 11 | 9453 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 12 | 9448 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 13 | 9452 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 14 | 9453 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 15 | 9450 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 16 | 9452 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 17 | 9449 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 18 | 9450 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 19 | 9449 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 20 | 9451 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 21 | 9451 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 22 | 9448 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 23 | 9451 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 24 | 9448 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 25 | 9452 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 26 | 9453 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 27 | 9451 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 28 | 9454 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 29 | 9448 | -86.83 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 30 | 9455 | -86.82 | 100.036 | 22.638565 | 22.492843 | 23.230582 | 58.042967 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 1 | 433.3 | -3.824 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 2 | 434.3 | -3.825 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 3 | 433.9 | -3.824 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 4 | 433.2 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 5 | 434.5 | -3.823 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 6 | 433.1 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 7 | 434.5 | -3.823 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 8 | 432.8 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 9 | 433.6 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 10 | 433.8 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 11 | 433.7 | -3.82 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 12 | 434.1 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 13 | 433.5 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 14 | 435.1 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 15 | 433.2 | -3.824 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 16 | 434.7 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 17 | 433.5 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 18 | 433.6 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 19 | 435.1 | -3.82 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 20 | 433.6 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 21 | 435.7 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 22 | 433.6 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 23 | 434.1 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 24 | 434.2 | -3.824 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 25 | 433.9 | -3.823 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 26 | 434.9 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 27 | 433.5 | -3.821 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 28 | 434.4 | -3.822 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 29 | 433.7 | -3.823 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 30 | 434.4 | -3.824 | 100.033 | 22.636151 | 22.481802 | 23.23695 | 58.027685 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 1 | 921.6 | -7.6 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 2 | 921.1 | -7.597 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 3 | 922.1 | -7.599 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 4 | 920.3 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 5 | 921.4 | -7.603 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 6 | 920.6 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 7 | 921.1 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 8 | 922.9 | -7.605 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 9 | 920.7 | -7.605 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 10 | 922.8 | -7.606 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 11 | 920.5 | -7.605 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 12 | 921.7 | -7.604 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 13 | 921.8 | -7.604 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 14 | 921.6 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 15 | 922.1 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 16 | 921.7 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 17 | 922.1 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 18 | 920.9 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 19 | 921.4 | -7.603 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 20 | 920.7 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 21 | 921.3 | -7.6 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 22 | 923.3 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 23 | 921.1 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 24 | 922.7 | -7.603 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 25 | 921.1 | -7.603 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 26 | 922.6 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 27 | 921.9 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 28 | 921.8 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 29 | 922.3 | -7.601 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 30 | 921.4 | -7.602 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 1 | 1509 | -12.02 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 2 | 1508 | -12.02 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 3 | 1509 | -12.02 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 4 | 1510 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 5 | 1509 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 6 | 1510 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 7 | 1508 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 8 | 1508 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 9 | 1509 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 10 | 1509 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 11 | 1508 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 12 | 1508 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 13 | 1510 | -12.01 | 100.034 | 22.641968 | 22.481304 | 23.238688 | 58.005232 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 14 | 1508 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 15 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 16 | 1508 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 17 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 18 | 1511 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 19 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 20 | 1510 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 21 | 1508 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 22 | 1508 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 23 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 24 | 1509 | -12.02 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 25 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 26 | 1508 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 27 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 28 | 1508 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 29 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 30 | 1509 | -12.01 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 1 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 2 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 3 | 2148 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 4 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 5 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 6 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 7 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 8 | 2148 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 9 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 10 | 2148 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 11 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 12 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 13 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 14 | 2152 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 15 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 16 | 2152 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 17 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 18 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 19 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 20 | 2151 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 21 | 2152 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 22 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 23 | 2151 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 24 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 25 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 26 | 2152 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 27 | 2150 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 28 | 2152 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 29 | 2149 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 30 | 2151 | -16.88 | 100.032 | 22.641178 | 22.480354 | 23.239943 | 57.980245 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 1 | 2876 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 2 | 2877 | -22.49 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 3 | 2877 | -22.49 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 4 | 2876 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 5 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 6 | 2876 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 7 | 2878 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 8 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 9 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 10 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 11 | 2876 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 12 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 13 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 14 | 2878 | -22.47 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 15 | 2878 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 16 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 17 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 18 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 19 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 20 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 21 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 22 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 23 | 2877 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 24 | 2880 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 25 | 2876 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 26 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 27 | 2879 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 28 | 2878 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 29 | 2878 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 30 | 2878 | -22.48 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 1 | 3672 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 2 | 3670 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 3 | 3672 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 4 | 3677 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 5 | 3675 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 6 | 3678 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 7 | 3674 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 8 | 3675 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 9 | 3675 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 10 | 3675 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 11 | 3675 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 12 | 3675 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 13 | 3676 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 14 | 3674 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 15 | 3676 | -28.72 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 16 | 3674 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 17 | 3676 | -28.71 | 100.034 | 22.638064 | 22.476731 | 23.24104 | 57.988739 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 18 | 3677 | -28.71 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 19 | 3675 | -28.71 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 20 | 3679 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 21 | 3675 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 22 | 3675 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 23 | 3675 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 24 | 3675 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 25 | 3677 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 26 | 3675 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 27 | 3676 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 28 | 3674 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 29 | 3676 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 30 | 3675 | -28.72 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 1 | 805.6 | -6.52 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 2 | 807.1 | -6.519 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 3 | 805.7 | -6.513 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 4 | 806.3 | -6.513 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 5 | 807.4 | -6.513 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 6 | 805.8 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 7 | 807.1 | -6.508 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 8 | 805.9 | -6.508 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 9 | 806.5 | -6.509 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 10 | 806.2 | -6.508 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 11 | 806.2 | -6.509 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 12 | 806.8 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 13 | 805.5 | -6.509 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 14 | 806.8 | -6.512 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 15 | 805.5 | -6.513 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 16 | 806.7 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 17 | 805.5 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 18 | 806.1 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 19 | 807.7 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 20 | 805.6 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 21 | 807.8 | -6.511 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 22 | 806.2 | -6.508 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 23 | 806.4 | -6.508 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 24 | 806.5 | -6.509 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 25 | 806.6 | -6.505 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 26 | 807.1 | -6.509 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 27 | 805.5 | -6.506 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 28 | 806.2 | -6.507 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 29 | 805.8 | -6.506 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 30 | 806.3 | -6.504 | 100.035 | 22.638098 | 22.476016 | 23.242458 | 57.980925 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 1 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 2 | 1938 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 3 | 1939 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 4 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 5 | 1938 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 6 | 1936 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 7 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 8 | 1939 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 9 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 10 | 1939 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 11 | 1936 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 12 | 1939 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 13 | 1938 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 14 | 1938 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 15 | 1938 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 16 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 17 | 1938 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 18 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 19 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 20 | 1937 | -14.92 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 21 | 1938 | -14.92 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 22 | 1939 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 23 | 1936 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 24 | 1939 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 25 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 26 | 1938 | -14.92 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 27 | 1938 | -14.92 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 28 | 1937 | -14.92 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 29 | 1939 | -14.92 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 30 | 1937 | -14.91 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 1 | 4169 | -32.18 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 2 | 4166 | -32.17 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 3 | 4167 | -32.17 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 4 | 4169 | -32.17 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 5 | 4168 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 6 | 4169 | -32.17 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 7 | 4168 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 8 | 4168 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 9 | 4167 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 10 | 4168 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 11 | 4167 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 12 | 4168 | -32.16 | 100.036 | 22.638492 | 22.473381 | 23.244258 | 57.991916 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 13 | 4171 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 14 | 4167 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 15 | 4170 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 16 | 4167 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 17 | 4169 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 18 | 4170 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 19 | 4168 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 20 | 4171 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 21 | 4167 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 22 | 4170 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 23 | 4167 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 24 | 4168 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 25 | 4168 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 26 | 4168 | -32.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 27 | 4169 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 28 | 4167 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 29 | 4169 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 30 | 4167 | -32.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 1 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 2 | 5906 | -46.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 3 | 5904 | -46.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 4 | 5906 | -46.16 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 5 | 5903 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 6 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 7 | 5904 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 8 | 5905 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 9 | 5907 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 10 | 5904 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 11 | 5910 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 12 | 5904 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 13 | 5905 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 14 | 5906 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 15 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 16 | 5908 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 17 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 18 | 5906 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 19 | 5904 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 20 | 5906 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 21 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 22 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 23 | 5908 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 24 | 5904 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 25 | 5908 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 26 | 5904 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 27 | 5906 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 28 | 5907 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 29 | 5905 | -46.17 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
120 | 19.9 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 30 | 5906 | -46.18 | 100.035 | 22.637333 | 22.471507 | 23.245276 | 57.988799 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 1 | 7781 | -61.85 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 2 | 7783 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 3 | 7783 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 4 | 7784 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 5 | 7785 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 6 | 7782 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 7 | 7785 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 8 | 7781 | -61.67 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 9 | 7784 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 10 | 7782 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 11 | 7782 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 12 | 7784 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 13 | 7783 | -61.62 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 14 | 7785 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 15 | 7781 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 16 | 7784 | -62.09 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 17 | 7782 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 18 | 7782 | -62 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 19 | 7787 | -61.83 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 20 | 7782 | -62.03 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 21 | 7786 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 22 | 7782 | -62.01 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 23 | 7783 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 24 | 7783 | -62.02 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 25 | 7783 | -62.07 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 26 | 7784 | -61.85 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 27 | 7783 | -61.85 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 28 | 7784 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 29 | 7783 | -61.61 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
140 | 19.9 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 30 | 7783 | -61.84 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 1 | 1465 | -10.69 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 2 | 1463 | -10.69 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 3 | 1463 | -10.69 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 4 | 1463 | -10.69 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 5 | 1463 | -10.69 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 6 | 1463 | -10.69 | 100.035 | 22.637573 | 22.46982 | 23.246256 | 57.988185 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 7 | 1462 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 8 | 1463 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 9 | 1461 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 10 | 1462 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 11 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 12 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 13 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 14 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 15 | 1464 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 16 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 17 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 18 | 1462 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 19 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 20 | 1463 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 21 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 22 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 23 | 1462 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 24 | 1462 | -10.69 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 25 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 26 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 27 | 1464 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 28 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 29 | 1463 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 30 | 1462 | -10.68 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 1 | 3354 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 2 | 3354 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 3 | 3354 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 4 | 3357 | -24.9 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 5 | 3353 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 6 | 3357 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 7 | 3356 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 8 | 3355 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 9 | 3356 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 10 | 3354 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 11 | 3356 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 12 | 3355 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 13 | 3356 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 14 | 3353 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 15 | 3355 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 16 | 3354 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 17 | 3355 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 18 | 3356 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 19 | 3353 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 20 | 3356 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 21 | 3354 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 22 | 3356 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 23 | 3356 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 24 | 3355 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 25 | 3356 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 26 | 3354 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 27 | 3355 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 28 | 3355 | -24.91 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 29 | 3354 | -24.92 | 100.035 | 22.638206 | 22.467192 | 23.247427 | 57.946628 | Set 1 on client chamber
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 30 | 3355 | -24.92 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 1 | 4872 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 2 | 4875 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 3 | 4871 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 4 | 4871 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 5 | 4872 | -36.92 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 6 | 4872 | -36.92 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 7 | 4873 | -36.92 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 8 | 4871 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 9 | 4875 | -36.92 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 10 | 4872 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 11 | 4874 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 12 | 4873 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 13 | 4874 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 14 | 4875 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 15 | 4873 | -36.93 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 16 | 4875 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 17 | 4872 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 18 | 4872 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 19 | 4873 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 20 | 4874 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 21 | 4874 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 22 | 4872 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 23 | 4875 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 24 | 4873 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 25 | 4874 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 26 | 4873 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 27 | 4874 | -36.95 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 28 | 4876 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 29 | 4874 | -36.95 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
120 | 19.9 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 30 | 4874 | -36.94 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 1 | 6539 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 2 | 6541 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 3 | 6542 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 4 | 6540 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 5 | 6542 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 6 | 6538 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 7 | 6543 | -50.6 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 8 | 6541 | -50.59 | 100.033 | 22.639163 | 22.467118 | 23.248457 | 57.938689 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 9 | 6541 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 10 | 6542 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 11 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 12 | 6542 | -50.6 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 13 | 6539 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 14 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 15 | 6539 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 16 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 17 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 18 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 19 | 6543 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 20 | 6539 | -50.6 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 21 | 6542 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 22 | 6539 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 23 | 6540 | -50.58 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 24 | 6543 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 25 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 26 | 6542 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 27 | 6539 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 28 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 29 | 6540 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 30 | 6541 | -50.59 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 1 | 7435 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 2 | 7434 | -58.1 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 3 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 4 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 5 | 7435 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 6 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 7 | 7433 | -58.12 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 8 | 7436 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 9 | 7432 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 10 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 11 | 7432 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 12 | 7433 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 13 | 7435 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 14 | 7433 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 15 | 7437 | -58.1 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 16 | 7432 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 17 | 7433 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 18 | 7432 | -58.1 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 19 | 7434 | -58.1 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 20 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 21 | 7433 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 22 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 23 | 7433 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 24 | 7435 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 25 | 7433 | -58.1 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 26 | 7432 | -58.1 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 27 | 7437 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 28 | 7434 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 29 | 7436 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 30 | 7433 | -58.11 | 100.033 | 22.637847 | 22.465659 | 23.249645 | 57.951707 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 1 | 2352 | -16.5 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 2 | 2353 | -16.49 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 3 | 2351 | -16.49 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 4 | 2352 | -16.49 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 5 | 2351 | -16.49 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 6 | 2352 | -16.49 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 7 | 2352 | -16.49 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 8 | 2351 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 9 | 2353 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 10 | 2352 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 11 | 2354 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 12 | 2351 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 13 | 2352 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 14 | 2353 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 15 | 2352 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 16 | 2354 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 17 | 2352 | -16.47 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 18 | 2352 | -16.47 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 19 | 2351 | -16.47 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 20 | 2352 | -16.47 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 21 | 2352 | -16.47 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 22 | 2352 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 23 | 2353 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 24 | 2351 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 25 | 2352 | -16.47 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 26 | 2351 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 27 | 2352 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 28 | 2354 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 29 | 2352 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 30 | 2354 | -16.48 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 1 | 3556 | -25.67 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 2 | 3558 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 3 | 3556 | -25.67 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 4 | 3557 | -25.67 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 5 | 3559 | -25.67 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 6 | 3557 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 7 | 3558 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 8 | 3556 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 9 | 3557 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 10 | 3557 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 11 | 3556 | -25.66 | 100.034 | 22.637917 | 22.46345 | 23.250418 | 57.942748 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 12 | 3558 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 13 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 14 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 15 | 3556 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 16 | 3556 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 17 | 3559 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 18 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 19 | 3559 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 20 | 3556 | -25.67 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 21 | 3559 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 22 | 3558 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 23 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 24 | 3558 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 25 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 26 | 3558 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 27 | 3556 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 28 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 29 | 3557 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
120 | 19.9 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 30 | 3556 | -25.66 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 1 | 4914 | -36.42 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 2 | 4914 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 3 | 4913 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 4 | 4912 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 5 | 4913 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 6 | 4913 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 7 | 4914 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 8 | 4914 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 9 | 4913 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 10 | 4914 | -36.43 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 11 | 4912 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 12 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 13 | 4914 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 14 | 4911 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 15 | 4914 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 16 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 17 | 4914 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 18 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 19 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 20 | 4914 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 21 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 22 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 23 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 24 | 4913 | -36.45 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 25 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 26 | 4914 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 27 | 4915 | -36.45 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 28 | 4914 | -36.45 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 29 | 4914 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 30 | 4913 | -36.44 | 100.034 | 22.63701 | 22.463265 | 23.251865 | 57.908982 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 1 | 5651 | -42.45 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 2 | 5649 | -42.46 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 3 | 5650 | -42.46 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 4 | 5651 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 5 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 6 | 5653 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 7 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 8 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 9 | 5651 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 10 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 11 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 12 | 5649 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 13 | 5649 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 14 | 5649 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 15 | 5649 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 16 | 5651 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 17 | 5650 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 18 | 5652 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 19 | 5648 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 20 | 5650 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 21 | 5651 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 22 | 5651 | -42.48 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 23 | 5652 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 24 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 25 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 26 | 5649 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 27 | 5650 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 28 | 5653 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 29 | 5649 | -42.46 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 30 | 5653 | -42.47 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 1 | 7383 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 2 | 7382 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 3 | 7383 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 4 | 7383 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 5 | 7385 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 6 | 7384 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 7 | 7385 | -58.68 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 8 | 7384 | -58.68 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 9 | 7384 | -58.68 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 10 | 7383 | -58.68 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 11 | 7382 | -58.69 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 12 | 7385 | -58.68 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 13 | 7383 | -58.67 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 14 | 7385 | -58.67 | 100.033 | 22.636512 | 22.46529 | 23.252695 | 57.880676 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 15 | 7383 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 16 | 7383 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 17 | 7384 | -58.68 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 18 | 7384 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 19 | 7386 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 20 | 7382 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 21 | 7385 | -58.66 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 22 | 7384 | -58.66 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 23 | 7384 | -58.66 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 24 | 7385 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 25 | 7382 | -58.68 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 26 | 7382 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 27 | 7382 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 28 | 7383 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 29 | 7382 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 30 | 7384 | -58.67 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 1 | 1954 | -15.26 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 2 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 3 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 4 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 5 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 6 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 7 | 1953 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 8 | 1953 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 9 | 1953 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 10 | 1954 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 11 | 1954 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 12 | 1954 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 13 | 1954 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 14 | 1953 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 15 | 1954 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 16 | 1953 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 17 | 1954 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 18 | 1955 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 19 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 20 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 21 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 22 | 1955 | -15.24 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 23 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 24 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 25 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 26 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 27 | 1955 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 28 | 1953 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 29 | 1954 | -15.25 | 100.035 | 22.641433 | 22.467952 | 23.253931 | 57.843679 | Set 1 on client chamber
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 30 | 1954 | -15.25 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 1 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 2 | 2943 | -23.83 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 3 | 2942 | -23.83 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 4 | 2943 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 5 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 6 | 2943 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 7 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 8 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 9 | 2944 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 10 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 11 | 2945 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 12 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 13 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 14 | 2943 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 15 | 2943 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 16 | 2944 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 17 | 2943 | -23.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 18 | 2943 | -23.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 19 | 2942 | -23.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 20 | 2943 | -23.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 21 | 2942 | -23.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 22 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 23 | 2944 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 24 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 25 | 2944 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 26 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 27 | 2943 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 28 | 2943 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 29 | 2942 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 30 | 2944 | -23.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 1 | 3503 | -28.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 2 | 3504 | -28.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 3 | 3503 | -28.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 4 | 3503 | -28.84 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 5 | 3504 | -28.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 6 | 3503 | -28.85 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 7 | 3504 | -28.86 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 8 | 3503 | -28.86 | 100.035 | 22.642694 | 22.471313 | 23.25445 | 57.842469 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 9 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 10 | 3502 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 11 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 12 | 3506 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 13 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 14 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 15 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 16 | 3506 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 17 | 3505 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 18 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 19 | 3504 | -28.87 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 20 | 3504 | -28.87 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 21 | 3505 | -28.87 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 22 | 3503 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 23 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 24 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 25 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 26 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 27 | 3504 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 28 | 3505 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 29 | 3503 | -28.86 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 30 | 3505 | -28.87 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 1 | 5160 | -45.27 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 2 | 5160 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 3 | 5161 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 4 | 5160 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 5 | 5161 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 6 | 5159 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 7 | 5160 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 8 | 5163 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 9 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 10 | 5162 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 11 | 5158 | -45.26 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 12 | 5161 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 13 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 14 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 15 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 16 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 17 | 5161 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 18 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 19 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 20 | 5160 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 21 | 5159 | -45.24 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 22 | 5161 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 23 | 5159 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 24 | 5163 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 25 | 5159 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 26 | 5161 | -45.25 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 27 | 5161 | -45.24 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 28 | 5159 | -45.24 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 29 | 5161 | -45.24 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 30 | 5159 | -45.24 | 100.035 | 22.64514 | 22.475338 | 23.254993 | 57.837564 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 1 | 6719 | -61.34 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 2 | 6715 | -61.34 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 3 | 6716 | -61.35 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 4 | 6717 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 5 | 6716 | -61.35 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 6 | 6718 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 7 | 6714 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 8 | 6715 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 9 | 6715 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 10 | 6715 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 11 | 6716 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 12 | 6715 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 13 | 6716 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 14 | 6716 | -61.38 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 15 | 6717 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 16 | 6716 | -61.53 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 17 | 6716 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 18 | 6717 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 19 | 6716 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 20 | 6718 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 21 | 6715 | -61.36 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 22 | 6715 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 23 | 6715 | -61.19 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 24 | 6716 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 25 | 6717 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 26 | 6715 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 27 | 6716 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 28 | 6714 | -61.38 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 29 | 6716 | -61.37 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 30 | 6714 | -61.38 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 1 | 1487 | -12.34 | 100.035 | 22.645253 | 22.479657 | 23.256444 | 57.813168 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 2 | 1488 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 3 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 4 | 1487 | -12.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 5 | 1487 | -12.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 6 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 7 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 8 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 9 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 10 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 11 | 1486 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 12 | 1486 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 13 | 1486 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 14 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 15 | 1486 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 16 | 1488 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 17 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 18 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 19 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 20 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 21 | 1488 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 22 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 23 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 24 | 1486 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 25 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 26 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 27 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 28 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 29 | 1487 | -12.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 30 | 1487 | -12.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 1 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 2 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 3 | 2171 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 4 | 2169 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 5 | 2171 | -18.34 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 6 | 2169 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 7 | 2171 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 8 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 9 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 10 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 11 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 12 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 13 | 2169 | -18.36 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 14 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 15 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 16 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 17 | 2171 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 18 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 19 | 2171 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 20 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 21 | 2171 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 22 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 23 | 2170 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 24 | 2171 | -18.35 | 100.033 | 22.648469 | 22.483093 | 23.257135 | 57.793009 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 25 | 2170 | -18.35 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 26 | 2171 | -18.35 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 27 | 2170 | -18.35 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 28 | 2170 | -18.35 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 29 | 2170 | -18.35 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 30 | 2170 | -18.35 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 1 | 3550 | -32.06 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 2 | 3550 | -32.06 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 3 | 3551 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 4 | 3549 | -32.06 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 5 | 3549 | -32.06 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 6 | 3550 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 7 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 8 | 3550 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 9 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 10 | 3551 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 11 | 3549 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 12 | 3550 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 13 | 3550 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 14 | 3549 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 15 | 3551 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 16 | 3549 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 17 | 3551 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 18 | 3549 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 19 | 3550 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 20 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 21 | 3550 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 22 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 23 | 3548 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 24 | 3550 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 25 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 26 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 27 | 3549 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 28 | 3549 | -32.08 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 29 | 3551 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 30 | 3549 | -32.07 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 1 | 4926 | -46.31 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 2 | 4924 | -46.31 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 3 | 4925 | -46.31 | 100.035 | 22.651482 | 22.486283 | 23.258051 | 57.746559 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 4 | 4926 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 5 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 6 | 4926 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 7 | 4925 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 8 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 9 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 10 | 4925 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 11 | 4925 | -46.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 12 | 4925 | -46.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 13 | 4926 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 14 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 15 | 4925 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 16 | 4924 | -46.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 17 | 4925 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 18 | 4926 | -46.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 19 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 20 | 4926 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 21 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 22 | 4925 | -46.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 23 | 4926 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 24 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 25 | 4926 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 26 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 27 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 28 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 29 | 4924 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 30 | 4925 | -46.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 1 | 5367 | -51.3 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 2 | 5367 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 3 | 5366 | -51.3 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 4 | 5368 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 5 | 5367 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 6 | 5367 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 7 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 8 | 5368 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 9 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 10 | 5367 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 11 | 5369 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 12 | 5367 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 13 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 14 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 15 | 5366 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 16 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 17 | 5368 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 18 | 5369 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 19 | 5367 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 20 | 5367 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 21 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 22 | 5367 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 23 | 5369 | -51.31 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 24 | 5367 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 25 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 26 | 5368 | -51.32 | 100.035 | 22.652508 | 22.488954 | 23.258794 | 57.732528 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 27 | 5368 | -51.32 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 28 | 5369 | -51.32 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 29 | 5368 | -51.32 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 30 | 5369 | -51.31 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 1 | 1068 | -9.185 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 2 | 1069 | -9.182 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 3 | 1069 | -9.178 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 4 | 1068 | -9.18 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 5 | 1069 | -9.178 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 6 | 1067 | -9.175 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 7 | 1068 | -9.175 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 8 | 1068 | -9.171 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 9 | 1067 | -9.172 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 10 | 1068 | -9.172 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 11 | 1068 | -9.172 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 12 | 1069 | -9.171 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 13 | 1068 | -9.167 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 14 | 1069 | -9.171 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 15 | 1068 | -9.17 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 16 | 1069 | -9.171 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 17 | 1068 | -9.167 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 18 | 1068 | -9.17 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 19 | 1068 | -9.169 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 20 | 1068 | -9.171 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 21 | 1068 | -9.169 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 22 | 1068 | -9.168 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 23 | 1068 | -9.167 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 24 | 1069 | -9.165 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 25 | 1068 | -9.163 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 26 | 1069 | -9.164 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 27 | 1067 | -9.162 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 28 | 1069 | -9.165 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 29 | 1068 | -9.166 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 30 | 1068 | -9.164 | 100.033 | 22.654192 | 22.491744 | 23.259657 | 57.717645 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 1 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 2 | 2592 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 3 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 4 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 5 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 6 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 7 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 8 | 2594 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 9 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 10 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 11 | 2592 | -23.81 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 12 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 13 | 2594 | -23.81 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 14 | 2592 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 15 | 2594 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 16 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 17 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 18 | 2592 | -23.81 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 19 | 2592 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 20 | 2592 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 21 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 22 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 23 | 2592 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 24 | 2592 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 25 | 2592 | -23.83 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 26 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 27 | 2594 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 28 | 2593 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 29 | 2594 | -23.82 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 30 | 2593 | -23.81 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 1 | 3797 | -36.24 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 2 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 3 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 4 | 3799 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 5 | 3798 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 6 | 3798 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 7 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 8 | 3798 | -36.22 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 9 | 3798 | -36.22 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 10 | 3796 | -36.22 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 11 | 3799 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 12 | 3798 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 13 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 14 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 15 | 3798 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 16 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 17 | 3798 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 18 | 3799 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 19 | 3798 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 20 | 3799 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 21 | 3797 | -36.23 | 100.031 | 22.657699 | 22.496365 | 23.260847 | 57.690841 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 22 | 3798 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 23 | 3797 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 24 | 3798 | -36.22 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 25 | 3799 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 26 | 3798 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 27 | 3797 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 28 | 3797 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 29 | 3797 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 30 | 3798 | -36.23 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 1 | 4244 | -41.11 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 2 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 3 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 4 | 4245 | -41.13 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 5 | 4245 | -41.13 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 6 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 7 | 4246 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 8 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 9 | 4244 | -41.11 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 10 | 4245 | -41.11 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 11 | 4245 | -41.11 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 12 | 4244 | -41.11 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 13 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 14 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 15 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 16 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 17 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 18 | 4245 | -41.13 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 19 | 4246 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 20 | 4244 | -41.13 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 21 | 4246 | -41.13 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 22 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 23 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 24 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 25 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 26 | 4245 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 27 | 4243 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 28 | 4246 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 29 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 30 | 4244 | -41.12 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 1 | 5048 | -49.32 | 100.029 | 22.659701 | 22.500299 | 23.2617 | 57.644981 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 2 | 5048 | -49.32 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 3 | 5048 | -49.31 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 4 | 5047 | -49.31 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 5 | 5049 | -49.3 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 6 | 5048 | -49.3 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 7 | 5048 | -49.3 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 8 | 5049 | -49.3 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 9 | 5048 | -49.3 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 10 | 5049 | -49.29 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 11 | 5048 | -49.29 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 12 | 5049 | -49.29 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 13 | 5048 | -49.29 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 14 | 5048 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 15 | 5049 | -49.29 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 16 | 5048 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 17 | 5049 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 18 | 5047 | -49.29 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 19 | 5048 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 20 | 5048 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 21 | 5049 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 22 | 5050 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 23 | 5049 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 24 | 5050 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 25 | 5048 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 26 | 5050 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 27 | 5049 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 28 | 5049 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 29 | 5049 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 30 | 5048 | -49.28 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 1 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 2 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 3 | 1441 | -13.27 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 4 | 1442 | -13.27 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 5 | 1441 | -13.27 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 6 | 1441 | -13.27 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 7 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 8 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 9 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 10 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 11 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 12 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 13 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 14 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 15 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 16 | 1441 | -13.26 | 100.028 | 22.662018 | 22.502305 | 23.262525 | 57.62555 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 17 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 18 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 19 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 20 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 21 | 1440 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 22 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 23 | 1442 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 24 | 1440 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 25 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 26 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 27 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 28 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 29 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 30 | 1441 | -13.26 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 1 | 2506 | -23.8 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 2 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 3 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 4 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 5 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 6 | 2506 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 7 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 8 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 9 | 2508 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 10 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 11 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 12 | 2505 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 13 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 14 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 15 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 16 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 17 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 18 | 2506 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 19 | 2506 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 20 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 21 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 22 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 23 | 2507 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 24 | 2506 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 25 | 2506 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 26 | 2505 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 27 | 2505 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 28 | 2507 | -23.82 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 29 | 2505 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 30 | 2507 | -23.81 | 100.028 | 22.660856 | 22.503038 | 23.263308 | 57.631785 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 1 | 2919 | -28.05 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 2 | 2920 | -28.05 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 3 | 2920 | -28.05 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 4 | 2919 | -28.05 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 5 | 2920 | -28.05 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 6 | 2920 | -28.06 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 7 | 2920 | -28.06 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 8 | 2920 | -28.07 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 9 | 2919 | -28.07 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 10 | 2920 | -28.07 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 11 | 2920 | -28.06 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 12 | 2920 | -28.07 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 13 | 2919 | -28.07 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 14 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 15 | 2921 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 16 | 2920 | -28.07 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 17 | 2922 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 18 | 2919 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 19 | 2921 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 20 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 21 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 22 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 23 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 24 | 2921 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 25 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 26 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 27 | 2919 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 28 | 2921 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 29 | 2921 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 30 | 2920 | -28.08 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 1 | 3553 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 2 | 3552 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 3 | 3554 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 4 | 3554 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 5 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 6 | 3554 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 7 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 8 | 3552 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 9 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 10 | 3552 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 11 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 12 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 13 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 14 | 3552 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 15 | 3553 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 16 | 3552 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 17 | 3552 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 18 | 3553 | -34.39 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 19 | 3553 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 20 | 3553 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 21 | 3553 | -34.4 | 100.031 | 22.665135 | 22.504644 | 23.264215 | 57.553703 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 22 | 3553 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 23 | 3553 | -34.39 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 24 | 3554 | -34.39 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 25 | 3553 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 26 | 3552 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 27 | 3553 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 28 | 3553 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 29 | 3553 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 30 | 3553 | -34.4 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 1 | 3418 | -33.26 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 2 | 3420 | -33.26 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 3 | 3418 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 4 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 5 | 3418 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 6 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 7 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 8 | 3419 | -33.26 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 9 | 3419 | -33.26 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 10 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 11 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 12 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 13 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 14 | 3420 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 15 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 16 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 17 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 18 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 19 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 20 | 3419 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 21 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 22 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 23 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 24 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 25 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 26 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 27 | 3418 | -33.27 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 28 | 3420 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 29 | 3419 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 30 | 3420 | -33.28 | 100.03 | 22.665262 | 22.506763 | 23.2652 | 57.519239 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 1 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 2 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 3 | 1648 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 4 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 5 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 6 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 7 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 8 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 9 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 10 | 1648 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 11 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 12 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 13 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 14 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 15 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 16 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 17 | 1647 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 18 | 1646 | -15.02 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 19 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 20 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 21 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 22 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 23 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 24 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 25 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 26 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 27 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 28 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 29 | 1647 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 30 | 1646 | -15.01 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 1 | 1989 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 2 | 1988 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 3 | 1988 | -18.31 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 4 | 1989 | -18.31 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 5 | 1988 | -18.31 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 6 | 1989 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 7 | 1988 | -18.31 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 8 | 1989 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 9 | 1989 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 10 | 1989 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 11 | 1989 | -18.31 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 12 | 1989 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 13 | 1990 | -18.32 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 14 | 1988 | -18.31 | 100.028 | 22.665205 | 22.505922 | 23.266071 | 57.473892 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 15 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 16 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 17 | 1988 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 18 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 19 | 1988 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 20 | 1990 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 21 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 22 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 23 | 1988 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 24 | 1988 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 25 | 1988 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 26 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 27 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 28 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 29 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 30 | 1989 | -18.32 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 1 | 2474 | -22.9 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 2 | 2473 | -22.9 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 3 | 2474 | -22.9 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 4 | 2474 | -22.9 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 5 | 2473 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 6 | 2473 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 7 | 2473 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 8 | 2473 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 9 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 10 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 11 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 12 | 2473 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 13 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 14 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 15 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 16 | 2474 | -22.91 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 17 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 18 | 2475 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 19 | 2473 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 20 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 21 | 2473 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 22 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 23 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 24 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 25 | 2475 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 26 | 2473 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 27 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 28 | 2473 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 29 | 2473 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 30 | 2474 | -22.92 | 100.031 | 22.667619 | 22.504638 | 23.267192 | 57.49951 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 1 | 2425 | -22.6 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 2 | 2427 | -22.6 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 3 | 2426 | -22.6 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 4 | 2426 | -22.6 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 5 | 2426 | -22.6 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 6 | 2426 | -22.6 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 7 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 8 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 9 | 2427 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 10 | 2427 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 11 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 12 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 13 | 2427 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 14 | 2428 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 15 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 16 | 2427 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 17 | 2426 | -22.58 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 18 | 2426 | -22.58 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 19 | 2427 | -22.58 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 20 | 2426 | -22.58 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 21 | 2427 | -22.58 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 22 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 23 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 24 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 25 | 2427 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 26 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 27 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 28 | 2427 | -22.58 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 29 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 30 | 2426 | -22.59 | 100.031 | 22.669585 | 22.504011 | 23.267906 | 57.52272 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 1 | 3808 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 2 | 3806 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 3 | 3807 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 4 | 3805 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 5 | 3807 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 6 | 3807 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 7 | 3806 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 8 | 3806 | -36.39 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 9 | 3805 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 10 | 3806 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 11 | 3806 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 12 | 3807 | -36.38 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 13 | 3807 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 14 | 3805 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 15 | 3808 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 16 | 3806 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 17 | 3806 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 18 | 3808 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 19 | 3807 | -36.36 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 20 | 3808 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 21 | 3807 | -36.36 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 22 | 3807 | -36.36 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 23 | 3806 | -36.36 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 24 | 3805 | -36.36 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 25 | 3808 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 26 | 3806 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 27 | 3807 | -36.36 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 28 | 3806 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 29 | 3806 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 30 | 3807 | -36.37 | 100.036 | 22.672708 | 22.500301 | 23.269062 | 57.46693 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 1 | 2920 | -28.06 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 2 | 2920 | -28.06 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 3 | 2919 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 4 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 5 | 2921 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 6 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 7 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 8 | 2919 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 9 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 10 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 11 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 12 | 2921 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 13 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 14 | 2921 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 15 | 2919 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 16 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 17 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 18 | 2920 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 19 | 2921 | -28.05 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 20 | 2919 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 21 | 2921 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 22 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 23 | 2921 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 24 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 25 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 26 | 2921 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 27 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 28 | 2921 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 29 | 2920 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 30 | 2921 | -28.04 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 1 | 3553 | -34.29 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 2 | 3554 | -34.3 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 3 | 3554 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 4 | 3553 | -34.3 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 5 | 3554 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 6 | 3554 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 7 | 3553 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 8 | 3554 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 9 | 3553 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 10 | 3554 | -34.31 | 100.038 | 22.676566 | 22.49986 | 23.269883 | 57.492065 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 11 | 3554 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 12 | 3555 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 13 | 3553 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 14 | 3554 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 15 | 3554 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 16 | 3553 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 17 | 3554 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 18 | 3553 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 19 | 3553 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 20 | 3553 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 21 | 3554 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 22 | 3555 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 23 | 3554 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 24 | 3553 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 25 | 3553 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 26 | 3554 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 27 | 3554 | -34.3 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 28 | 3555 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 29 | 3555 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 30 | 3554 | -34.31 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 1 | 0.1 | -0.002 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 2 | 0.15 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 3 | 0.2 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 4 | 0.3 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 5 | 0.55 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 6 | 0.25 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 7 | 0.05 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 8 | 0 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 9 | 0.2 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 10 | 0.15 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 11 | 0 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 12 | 0.2 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 13 | 0 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 14 | 0.15 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 15 | 0 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 16 | 0 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 17 | 0 | 0 | 100.038 | 22.676354 | 22.498172 | 23.271274 | 57.475626 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 18 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 19 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 20 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 21 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 22 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 23 | 0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 24 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 25 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 26 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 27 | 0.4 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 28 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 29 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 30 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 31 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 32 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 33 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 34 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 35 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 36 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 37 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 38 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 39 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 40 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 41 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 42 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 43 | 0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 44 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 45 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 46 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 47 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 48 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 49 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 50 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 51 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 52 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 53 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 54 | 0.4 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 55 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 56 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 57 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 58 | 0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 59 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 60 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 61 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 62 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 63 | 0.3 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 64 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 65 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 66 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 67 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 68 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 69 | -0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 70 | 0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 71 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 72 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 73 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 74 | 0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 75 | 0.25 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 76 | -0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 77 | 0.05 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 78 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 79 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 80 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 81 | -0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 82 | 0 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 83 | 0.1 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 84 | 0.2 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 85 | 0.15 | 0 | 100.038 | 22.676857 | 22.493188 | 23.272112 | 57.455939 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 86 | 0 | 0 | 100.04 | 22.675012 | 22.487107 | 23.27285 | 57.442629 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 87 | 0 | 0 | 100.04 | 22.675012 | 22.487107 | 23.27285 | 57.442629 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 88 | 0.3 | 0 | 100.04 | 22.675012 | 22.487107 | 23.27285 | 57.442629 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 89 | 0.3 | 0 | 100.04 | 22.675012 | 22.487107 | 23.27285 | 57.442629 | Set 1 on client chamber
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 90 | 0 | 0 | 100.04 | 22.675012 | 22.487107 | 23.27285 | 57.442629 | Set 1 on client chamber

## lab
[COMET X-RAY MEASUREMENT]
Filename |  | C:\CRData\12Feb2021-3-24P.csv
Date |  | 12/02/2021
Chamber |  | MEFAC
Description |  | Standard ratio measurement
Software |  | S:\Medical_Rad\Radiotherapy\Ionizing Radiation\IRS Inhouse Software\Twin Weblines LEX and MEX\Twin Webline Comet v8_0-PRODUCTION.vi
Backgrounds |  | 90
Measurements |  | 30
Trolley (mm) |  | 1090
SCD (mm) |  | 1000
Aperture wheel |  | 2 CM
Comment |  | MEFAC run
Monitor electrometer range |  | Medium
Monitor HV |  | 300
MEFAC-IC electrometer range |  | Medium
IC HV |  | 0
Client name |  | ClientA_Name | [6 empty] | Client identifier string |  |  | *this line may not be present in all files
Address 1 |  | 200 Street Name | [6 empty] | Address line 1 |  |  | *this line may not be present in all files
Address 2 |  | Suburb NSW 2020 | [6 empty] | Address line 2 |  |  | *this line may not be present in all files
Operator |  | Duncan Butler | [6 empty] | Who did the measurements |  |  | *this line may not be present in all files
CAL Number |  | CAL00001 | [6 empty] | ARPANSA Job ID |  |  | *this line may not be present in all files
[DATA]
kV | mA | BarCode | XraysOn | HVLFilter(mm) | Filter | FilterReady | HVLReady | N | Current1(pA) | Current2(pA) | P(kPa) | T(MC) | T(Air) | T(SC) | H(%) | Comment
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 1 | 0.25 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 2 | 0.15 | -0.15 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 3 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 4 | 0.05 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 5 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 6 | 0.15 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 7 | 0.15 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 8 | 0.15 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 9 | 0 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 10 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 11 | 0.3 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 12 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 13 | 0 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 14 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 15 | 0.15 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 16 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 17 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 18 | 0 | -0.15 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 19 | 0.2 | -0.15 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 20 | 0.2 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 21 | 0.2 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 22 | 0.15 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 23 | 0.2 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 24 | 0.05 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 25 | 0 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 26 | 0.25 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 27 | 0.1 | -0.1 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 28 | 0.05 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 29 | 0.3 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 30 | 0.15 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 31 | 0.35 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 32 | 0.15 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 33 | 0.3 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 34 | 0.2 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 35 | 0.2 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 36 | 0.15 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 37 | 0.2 | -0.1 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 38 | 0.15 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 39 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 40 | 0.15 | -0.15 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 41 | 0.4 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 42 | 0.3 | -0.05 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 43 | 0 | -0.15 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 44 | 0.25 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 45 | 0.2 | 0 | 100.06 | 22.486987 | 22.501162 | 22.964226 | 59.177399 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 46 | 0.2 | -0.05 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 47 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 48 | 0.3 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 49 | 0.3 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 50 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 51 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 52 | 0.25 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 53 | 0.1 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 54 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 55 | 0.1 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 56 | 0.2 | -0.05 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 57 | 0.25 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 58 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 59 | 0.05 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 60 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 61 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 62 | 0.25 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 63 | 0 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 64 | 0.15 | -0.05 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 65 | 0.15 | -0.05 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 66 | 0.25 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 67 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 68 | 0.1 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 69 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 70 | 0 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 71 | 0.1 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 72 | 0.25 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 73 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 74 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 75 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 76 | 0.05 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 77 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 78 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 79 | 0 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 80 | 0 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 81 | 0.05 | -0.05 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 82 | 0.2 | -0.1 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 83 | 0 | -0.1 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 84 | 0.15 | -0.1 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 85 | 0.15 | -0.05 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 86 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 87 | 0.05 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 88 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 89 | 0.2 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 90 | 0.15 | 0 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 1 | 3712 | -332 | 100.06 | 22.48655 | 22.501373 | 22.969547 | 59.431607 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 2 | 3705 | -331.3 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 3 | 3713 | -331.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 4 | 3707 | -331.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 5 | 3707 | -331.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 6 | 3708 | -331.1 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 7 | 3706 | -331.4 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 8 | 3709 | -331.3 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 9 | 3704 | -331.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 10 | 3709 | -331.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 11 | 3705 | -331.4 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 12 | 3708 | -331.8 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 13 | 3706 | -331.2 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 14 | 3706 | -332 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 15 | 3713 | -332.3 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 16 | 3705 | -331.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 17 | 3711 | -331.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 18 | 3706 | -331.2 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 19 | 3707 | -332.3 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 20 | 3710 | -331.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 21 | 3708 | -331.8 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 22 | 3712 | -331.6 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 23 | 3710 | -331.8 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 24 | 3710 | -332.2 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 25 | 3707 | -331.6 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 26 | 3709 | -331.8 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 27 | 3709 | -331.3 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 28 | 3709 | -331.8 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 29 | 3713 | -331.9 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
40 | 20 | NXB | TRUE | 0 | NXJ40 | FALSE | TRUE | 30 | 3707 | -331.9 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 1 | 5373 | -467 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 2 | 5366 | -466.6 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 3 | 5370 | -467 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 4 | 5372 | -466 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 5 | 5369 | -466.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 6 | 5373 | -466.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 7 | 5369 | -466.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 8 | 5375 | -466.4 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 9 | 5369 | -466.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 10 | 5370 | -466.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 11 | 5372 | -466.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 12 | 5371 | -466.9 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 13 | 5372 | -466.6 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 14 | 5369 | -466.7 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 15 | 5373 | -466.9 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 16 | 5368 | -466.6 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 17 | 5371 | -467.4 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 18 | 5368 | -466.4 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 19 | 5370 | -466.9 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 20 | 5375 | -467 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 21 | 5370 | -466.5 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 22 | 5375 | -467.1 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 23 | 5368 | -466.6 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 24 | 5370 | -467.2 | 100.06 | 22.485323 | 22.498312 | 22.972255 | 59.206522 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 25 | 5372 | -466.9 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 26 | 5371 | -466.5 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 27 | 5374 | -467.1 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 28 | 5370 | -466.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 29 | 5372 | -466.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXJ50 | FALSE | TRUE | 30 | 5369 | -466.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 1 | 6992 | -593.9 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 2 | 6991 | -594.2 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 3 | 6991 | -594.3 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 4 | 6998 | -594.1 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 5 | 6991 | -594.2 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 6 | 6995 | -595.2 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 7 | 6990 | -594 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 8 | 6992 | -594.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 9 | 6996 | -594.3 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 10 | 6991 | -594.1 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 11 | 6997 | -594.5 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 12 | 6990 | -593.9 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 13 | 6993 | -594.6 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 14 | 6992 | -594.6 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 15 | 6993 | -594.1 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 16 | 6993 | -594 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 17 | 6991 | -594 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 18 | 6996 | -594.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 19 | 6990 | -593.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 20 | 6994 | -594.7 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 21 | 6991 | -593.7 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 22 | 6993 | -594.4 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 23 | 6998 | -594.6 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 24 | 6992 | -594.1 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 25 | 6998 | -594.3 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 26 | 6992 | -594.4 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 27 | 6993 | -594.5 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 28 | 6994 | -593.9 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 29 | 6993 | -594.2 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
60 | 20 | NXB | TRUE | 0 | NXJ60 | FALSE | TRUE | 30 | 6996 | -594 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 1 | 8525 | -712.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 2 | 8531 | -712.8 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 3 | 8525 | -713 | 100.059 | 22.484879 | 22.495077 | 22.975444 | 59.096873 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 4 | 8527 | -713.6 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 5 | 8532 | -712.4 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 6 | 8528 | -713.3 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 7 | 8532 | -713.5 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 8 | 8527 | -713 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 9 | 8528 | -712.6 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 10 | 8526 | -712.8 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 11 | 8528 | -713.2 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 12 | 8529 | -712.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 13 | 8526 | -713.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 14 | 8532 | -713.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 15 | 8526 | -713.4 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 16 | 8531 | -713.2 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 17 | 8532 | -712.8 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 18 | 8528 | -712.9 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 19 | 8532 | -713.6 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 20 | 8526 | -713.1 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 21 | 8530 | -713.8 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 22 | 8528 | -712.9 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 23 | 8529 | -713.5 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 24 | 8533 | -713.4 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 25 | 8527 | -713.6 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 26 | 8531 | -713 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 27 | 8527 | -713 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 28 | 8530 | -713.3 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 29 | 8530 | -713.2 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXJ70 | FALSE | TRUE | 30 | 8528 | -713.6 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 1 | 10050 | -829.5 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 2 | 10050 | -829.5 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 3 | 10050 | -829.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 4 | 10050 | -830 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 5 | 10050 | -829.6 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 6 | 10050 | -828.9 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 7 | 10050 | -830.3 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 8 | 10050 | -829.4 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 9 | 10050 | -829.9 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 10 | 10060 | -829.5 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 11 | 10050 | -829.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 12 | 10060 | -829.9 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 13 | 10050 | -829.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 14 | 10050 | -830.3 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 15 | 10060 | -830.1 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 16 | 10050 | -829.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 17 | 10060 | -829.5 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 18 | 10050 | -830.2 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 19 | 10060 | -829.4 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 20 | 10050 | -829.7 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 21 | 10050 | -829.8 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 22 | 10050 | -830.1 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 23 | 10050 | -830.2 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 24 | 10060 | -830.3 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 25 | 10050 | -830.1 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 26 | 10060 | -830.1 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 27 | 10050 | -829.8 | 100.054 | 22.484085 | 22.489103 | 22.980507 | 59.009898 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 28 | 10050 | -829.6 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 29 | 10060 | -829.3 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
80 | 20 | NXB | TRUE | 0 | NXJ80 | FALSE | TRUE | 30 | 10050 | -829.9 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 1 | 11560 | -943.8 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 2 | 11550 | -944 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 3 | 11560 | -943.7 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 4 | 11550 | -944.1 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 5 | 11560 | -943.9 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 6 | 11560 | -943.3 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 7 | 11560 | -943.7 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 8 | 11560 | -944.1 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 9 | 11560 | -944 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 10 | 11560 | -943.7 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 11 | 11550 | -943.8 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 12 | 11560 | -944.1 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 13 | 11550 | -943.7 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 14 | 11560 | -944 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 15 | 11560 | -943.6 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 16 | 11550 | -944.3 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 17 | 11560 | -944.2 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 18 | 11560 | -944 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 19 | 11560 | -943.6 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 20 | 11560 | -943.5 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 21 | 11560 | -943.4 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 22 | 11560 | -943.7 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 23 | 11550 | -943.9 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 24 | 11560 | -944.1 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 25 | 11560 | -943.8 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 26 | 11560 | -943.8 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 27 | 11550 | -943.3 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 28 | 11550 | -944.2 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 29 | 11560 | -944 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
90 | 20 | NXB | TRUE | 0 | NXJ90 | FALSE | TRUE | 30 | 11560 | -943.5 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 1 | 13030 | -1056 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 2 | 13030 | -1056 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 3 | 13030 | -1055 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 4 | 13030 | -1056 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 5 | 13030 | -1055 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 6 | 13030 | -1056 | 100.054 | 22.485399 | 22.488302 | 22.983684 | 59.046549 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 7 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 8 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 9 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 10 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 11 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 12 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 13 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 14 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 15 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 16 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 17 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 18 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 19 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 20 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 21 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 22 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 23 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 24 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 25 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 26 | 13030 | -1056 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 27 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 28 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 29 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXJ100 | FALSE | TRUE | 30 | 13030 | -1055 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 1 | 1998 | -180.6 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 2 | 1999 | -180.7 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 3 | 1997 | -180.6 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 4 | 1999 | -180.8 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 5 | 2001 | -180.8 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 6 | 1997 | -180.8 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 7 | 2002 | -181 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 8 | 1997 | -180.5 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 9 | 1999 | -180.8 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 10 | 1998 | -180.7 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 11 | 1998 | -180.8 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 12 | 1999 | -180.9 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 13 | 1998 | -180.6 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 14 | 2000 | -180.7 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 15 | 1998 | -180.7 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 16 | 1999 | -180.8 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 17 | 1997 | -180.7 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 18 | 1998 | -181 | 100.054 | 22.486675 | 22.489853 | 22.988477 | 58.75665 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 19 | 2001 | -180.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 20 | 1997 | -180.7 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 21 | 2001 | -180.7 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 22 | 1997 | -180.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 23 | 2000 | -180.9 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 24 | 2000 | -180.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 25 | 1999 | -180.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 26 | 2000 | -180.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 27 | 1997 | -180.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 28 | 2001 | -180.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 29 | 1997 | -180.7 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
40 | 20 | NXK | TRUE | 0 | NXK40 | TRUE | TRUE | 30 | 1998 | -180.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 1 | 3173 | -273.4 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 2 | 3173 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 3 | 3174 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 4 | 3172 | -273.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 5 | 3176 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 6 | 3173 | -273.4 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 7 | 3175 | -273.9 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 8 | 3171 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 9 | 3173 | -273.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 10 | 3176 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 11 | 3174 | -273.2 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 12 | 3178 | -273.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 13 | 3174 | -273.2 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 14 | 3174 | -273.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 15 | 3173 | -273.7 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 16 | 3174 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 17 | 3174 | -273.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 18 | 3174 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 19 | 3176 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 20 | 3172 | -273.4 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 21 | 3175 | -273.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 22 | 3173 | -273.5 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 23 | 3174 | -273.7 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 24 | 3177 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 25 | 3174 | -273.6 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 26 | 3177 | -274.1 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 27 | 3173 | -273.4 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 28 | 3177 | -274 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 29 | 3174 | -273.3 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
50 | 20 | NXK | TRUE | 0 | NXK50 | TRUE | TRUE | 30 | 3174 | -273.8 | 100.051 | 22.489694 | 22.493064 | 22.991486 | 58.778822 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 1 | 4393 | -366 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 2 | 4393 | -366.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 3 | 4395 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 4 | 4392 | -366.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 5 | 4396 | -366.8 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 6 | 4393 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 7 | 4395 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 8 | 4394 | -366.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 9 | 4393 | -367 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 10 | 4398 | -366.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 11 | 4393 | -366.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 12 | 4398 | -366.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 13 | 4395 | -367.1 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 14 | 4395 | -366.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 15 | 4396 | -366.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 16 | 4394 | -366.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 17 | 4397 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 18 | 4393 | -366.4 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 19 | 4394 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 20 | 4394 | -366.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 21 | 4395 | -366.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 22 | 4394 | -366.8 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 23 | 4395 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 24 | 4397 | -366.8 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 25 | 4394 | -366.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 26 | 4396 | -366.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 27 | 4393 | -366.3 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 28 | 4394 | -366.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 29 | 4396 | -366.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
60 | 20 | NXK | TRUE | 0 | NXK60 | TRUE | TRUE | 30 | 4394 | -367 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 1 | 5608 | -457.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 2 | 5604 | -457.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 3 | 5609 | -458 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 4 | 5604 | -457.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 5 | 5605 | -457.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 6 | 5605 | -457.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 7 | 5605 | -457.4 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 8 | 5606 | -457.1 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 9 | 5605 | -457.3 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 10 | 5606 | -458.1 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 11 | 5604 | -457.4 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 12 | 5606 | -458.4 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 13 | 5606 | -457.1 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 14 | 5605 | -457.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 15 | 5608 | -457.9 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 16 | 5605 | -457.8 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 17 | 5609 | -457.7 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 18 | 5603 | -457.4 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 19 | 5605 | -457.6 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 20 | 5607 | -457.5 | 100.05 | 22.49377 | 22.499447 | 22.996198 | 58.877224 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 21 | 5606 | -457.9 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 22 | 5608 | -457.5 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 23 | 5604 | -457.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 24 | 5606 | -458.3 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 25 | 5604 | -457.9 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 26 | 5605 | -457.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 27 | 5605 | -457.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 28 | 5605 | -457.4 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 29 | 5608 | -457.5 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
70 | 20 | NXK | TRUE | 0 | NXK70 | TRUE | TRUE | 30 | 5604 | -457.9 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 1 | 6864 | -551.9 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 2 | 6860 | -551.8 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 3 | 6863 | -552 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 4 | 6865 | -551.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 5 | 6862 | -551.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 6 | 6866 | -551.8 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 7 | 6862 | -551.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 8 | 6863 | -551.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 9 | 6861 | -551.1 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 10 | 6863 | -551.8 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 11 | 6866 | -551.9 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 12 | 6863 | -551.1 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 13 | 6865 | -551.4 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 14 | 6861 | -551.4 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 15 | 6864 | -551.5 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 16 | 6865 | -551.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 17 | 6863 | -551.9 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 18 | 6865 | -551.8 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 19 | 6861 | -551.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 20 | 6862 | -551.8 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 21 | 6862 | -551.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 22 | 6863 | -551.4 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 23 | 6863 | -551.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 24 | 6863 | -551.3 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 25 | 6865 | -551.4 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 26 | 6861 | -551.5 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 27 | 6864 | -551.8 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 28 | 6861 | -550.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 29 | 6864 | -551.7 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
80 | 20 | NXK | TRUE | 0 | NXK80 | TRUE | TRUE | 30 | 6863 | -551.6 | 100.049 | 22.497594 | 22.502746 | 22.999274 | 58.794737 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 1 | 8146 | -647 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 2 | 8150 | -647.3 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 3 | 8145 | -647.1 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 4 | 8146 | -647.6 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 5 | 8145 | -647.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 6 | 8146 | -647.7 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 7 | 8147 | -647.2 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 8 | 8146 | -647 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 9 | 8148 | -647.3 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 10 | 8145 | -646.6 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 11 | 8148 | -647.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 12 | 8145 | -646.8 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 13 | 8147 | -647.1 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 14 | 8150 | -647.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 15 | 8145 | -647.1 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 16 | 8149 | -647.8 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 17 | 8145 | -647.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 18 | 8150 | -647.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 19 | 8146 | -647.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 20 | 8145 | -647 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 21 | 8145 | -647.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 22 | 8146 | -646.7 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 23 | 8148 | -648.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 24 | 8144 | -647.1 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 25 | 8148 | -647.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 26 | 8146 | -646.8 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 27 | 8147 | -647.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 28 | 8145 | -646.9 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 29 | 8145 | -647.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
90 | 20 | NXK | TRUE | 0 | NXK90 | TRUE | TRUE | 30 | 8149 | -647.2 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 1 | 9443 | -743.7 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 2 | 9449 | -744.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 3 | 9442 | -744.1 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 4 | 9444 | -744.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 5 | 9446 | -744.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 6 | 9445 | -743.3 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 7 | 9447 | -743.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 8 | 9444 | -743.7 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 9 | 9446 | -743.6 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 10 | 9444 | -743.5 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 11 | 9446 | -744.3 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 12 | 9445 | -743.6 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 13 | 9446 | -744 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 14 | 9450 | -743.6 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 15 | 9444 | -743.6 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 16 | 9448 | -745.2 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 17 | 9443 | -743.4 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 18 | 9447 | -744.2 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 19 | 9448 | -744 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 20 | 9445 | -744 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 21 | 9450 | -744.2 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 22 | 9445 | -743.8 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 23 | 9449 | -744.2 | 100.048 | 22.500912 | 22.504244 | 23.002145 | 58.62149 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 24 | 9445 | -744.5 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 25 | 9446 | -743.7 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 26 | 9444 | -743.8 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 27 | 9446 | -743.3 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 28 | 9446 | -744.2 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 29 | 9444 | -744.4 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
100 | 20 | NXK | TRUE | 0 | NXK100 | TRUE | TRUE | 30 | 9447 | -744 | 100.046 | 22.506867 | 22.50488 | 23.006558 | 58.498269 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 1 | 434.1 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 2 | 434.2 | -32.35 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 3 | 433.5 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 4 | 434.4 | -32.1 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 5 | 433.2 | -32.2 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 6 | 434.1 | -32.2 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 7 | 433 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 8 | 433.7 | -32.35 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 9 | 434.9 | -32.3 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 10 | 433.3 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 11 | 434.8 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 12 | 433.1 | -32.2 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 13 | 434.8 | -32.45 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 14 | 433.9 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 15 | 433.5 | -32.2 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 16 | 434.1 | -32.3 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 17 | 433.8 | -32.3 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 18 | 434.7 | -32.2 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 19 | 433.8 | -32.35 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 20 | 434.2 | -32.4 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 21 | 433.6 | -32.15 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 22 | 433.9 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 23 | 434 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 24 | 433.6 | -32.3 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 25 | 434.9 | -32.35 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 26 | 433.4 | -32.25 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 27 | 434.3 | -32.35 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 28 | 433.4 | -32.2 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 29 | 434 | -32.3 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
40 | 20 | NXA | TRUE | 0 | NXA40 | TRUE | TRUE | 30 | 435.1 | -32.35 | 100.047 | 22.513486 | 22.508912 | 23.00959 | 58.462393 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 1 | 920.7 | -63.95 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 2 | 921.8 | -64.2 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 3 | 920.5 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 4 | 921.2 | -64.1 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 5 | 920.7 | -64.05 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 6 | 920.7 | -64.3 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 7 | 920.4 | -64 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 8 | 921.1 | -63.9 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 9 | 921.2 | -64 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 10 | 920.2 | -64.05 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 11 | 921.5 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 12 | 919.8 | -63.75 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 13 | 921.4 | -64.1 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 14 | 920.7 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 15 | 920.7 | -64.05 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 16 | 922.7 | -64.05 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 17 | 920.7 | -63.95 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 18 | 922.1 | -64.1 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 19 | 920.6 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 20 | 921.6 | -64.3 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 21 | 922.2 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 22 | 922 | -64.1 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 23 | 922.8 | -64.05 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 24 | 921.4 | -64.1 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 25 | 921.7 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 26 | 921 | -64 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 27 | 921.4 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 28 | 921.5 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 29 | 921.5 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
50 | 20 | NXA | TRUE | 0 | NXA50 | TRUE | TRUE | 30 | 922.6 | -64.15 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 1 | 1508 | -101.2 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 2 | 1510 | -101.6 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 3 | 1507 | -101.6 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 4 | 1509 | -101.5 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 5 | 1509 | -101.4 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 6 | 1508 | -101.6 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 7 | 1510 | -101.5 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 8 | 1508 | -101.4 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 9 | 1509 | -101.4 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 10 | 1508 | -101.4 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 11 | 1509 | -101.2 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 12 | 1508 | -101.6 | 100.048 | 22.531776 | 22.518132 | 23.013933 | 58.472419 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 13 | 1508 | -101.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 14 | 1510 | -101.1 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 15 | 1508 | -101.2 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 16 | 1509 | -101.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 17 | 1508 | -101.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 18 | 1509 | -101.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 19 | 1510 | -101.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 20 | 1508 | -101.6 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 21 | 1510 | -101.6 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 22 | 1508 | -101.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 23 | 1510 | -101.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 24 | 1508 | -101.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 25 | 1508 | -101.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 26 | 1509 | -101.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 27 | 1509 | -101.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 28 | 1509 | -101.3 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 29 | 1507 | -101.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
60 | 20 | NXA | TRUE | 0 | NXA60 | TRUE | TRUE | 30 | 1509 | -101.9 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 1 | 2148 | -142.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 2 | 2150 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 3 | 2148 | -142.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 4 | 2149 | -142.8 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 5 | 2150 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 6 | 2148 | -142.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 7 | 2151 | -142.8 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 8 | 2148 | -142.6 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 9 | 2149 | -142.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 10 | 2149 | -142.3 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 11 | 2149 | -142.6 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 12 | 2150 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 13 | 2148 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 14 | 2149 | -142.8 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 15 | 2148 | -142.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 16 | 2149 | -143 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 17 | 2148 | -142.6 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 18 | 2148 | -142.9 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 19 | 2150 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 20 | 2148 | -142.6 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 21 | 2151 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 22 | 2148 | -142.5 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 23 | 2148 | -143 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 24 | 2150 | -143 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 25 | 2148 | -142.7 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 26 | 2151 | -142.9 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 27 | 2149 | -143.1 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 28 | 2150 | -143 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 29 | 2149 | -142.8 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
70 | 20 | NXA | TRUE | 0 | NXA70 | TRUE | TRUE | 30 | 2149 | -142.4 | 100.046 | 22.535891 | 22.524006 | 23.016986 | 58.409041 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 1 | 2874 | -190.1 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 2 | 2876 | -189.9 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 3 | 2875 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 4 | 2875 | -190.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 5 | 2877 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 6 | 2874 | -190.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 7 | 2877 | -190.7 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 8 | 2874 | -190.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 9 | 2876 | -190.6 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 10 | 2878 | -190 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 11 | 2875 | -190.1 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 12 | 2878 | -190.2 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 13 | 2875 | -190.2 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 14 | 2878 | -190.1 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 15 | 2876 | -190.2 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 16 | 2876 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 17 | 2877 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 18 | 2876 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 19 | 2877 | -190.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 20 | 2875 | -190.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 21 | 2877 | -190.6 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 22 | 2875 | -190.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 23 | 2876 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 24 | 2878 | -190 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 25 | 2875 | -190.1 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 26 | 2878 | -190.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 27 | 2876 | -190.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 28 | 2877 | -190.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 29 | 2877 | -190.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
80 | 20 | NXA | TRUE | 0 | NXA80 | TRUE | TRUE | 30 | 2876 | -190.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 1 | 3672 | -243.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 2 | 3672 | -243.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 3 | 3673 | -242.9 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 4 | 3671 | -243.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 5 | 3672 | -243.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 6 | 3671 | -243.5 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 7 | 3672 | -243.7 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 8 | 3672 | -243 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 9 | 3671 | -243.7 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 10 | 3674 | -243.6 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 11 | 3670 | -243.4 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 12 | 3674 | -243.7 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 13 | 3671 | -243.3 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 14 | 3672 | -243.8 | 100.044 | 22.539001 | 22.529277 | 23.020389 | 58.370918 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 15 | 3673 | -243.8 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 16 | 3672 | -243.2 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 17 | 3674 | -243.6 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 18 | 3671 | -243.5 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 19 | 3671 | -243.5 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 20 | 3671 | -243.4 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 21 | 3671 | -243.3 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 22 | 3671 | -243.2 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 23 | 3672 | -243.8 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 24 | 3673 | -243.4 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 25 | 3671 | -243.4 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 26 | 3672 | -244 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 27 | 3671 | -243.3 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 28 | 3672 | -243.8 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 29 | 3671 | -243.4 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
90 | 20 | NXA | TRUE | 0 | NXA90 | TRUE | TRUE | 30 | 3671 | -243.5 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 1 | 805 | -54.7 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 2 | 803.6 | -54.5 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 3 | 804.3 | -54.55 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 4 | 805.2 | -54.55 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 5 | 803.9 | -54.55 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 6 | 805.1 | -54.75 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 7 | 803.2 | -54.55 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 8 | 805.1 | -54.8 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 9 | 803.9 | -54.6 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 10 | 803.7 | -54.6 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 11 | 805.1 | -54.7 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 12 | 803.6 | -54.65 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 13 | 805.3 | -54.65 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 14 | 803.8 | -54.65 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 15 | 804.1 | -54.7 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 16 | 803.7 | -54.6 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 17 | 803.8 | -54.75 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 18 | 804.3 | -54.7 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 19 | 803.7 | -54.6 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 20 | 805.4 | -54.75 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 21 | 803.8 | -54.55 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 22 | 804.2 | -54.7 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 23 | 803.3 | -54.65 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 24 | 803.9 | -54.45 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 25 | 805 | -54.5 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 26 | 803.9 | -54.6 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 27 | 805.6 | -54.7 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 28 | 803.7 | -54.45 | 100.039 | 22.542742 | 22.534941 | 23.024665 | 58.361798 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 29 | 803.7 | -54.85 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
50 | 20 | NXB | TRUE | 0 | NXB50 | TRUE | TRUE | 30 | 803.4 | -54.6 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 1 | 1934 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 2 | 1934 | -126 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 3 | 1934 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 4 | 1935 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 5 | 1932 | -126.1 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 6 | 1935 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 7 | 1933 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 8 | 1934 | -126 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 9 | 1935 | -125.5 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 10 | 1934 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 11 | 1936 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 12 | 1933 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 13 | 1936 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 14 | 1933 | -126.1 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 15 | 1935 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 16 | 1934 | -126 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 17 | 1934 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 18 | 1935 | -125.7 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 19 | 1932 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 20 | 1934 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 21 | 1933 | -125.7 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 22 | 1935 | -125.7 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 23 | 1936 | -125.6 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 24 | 1934 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 25 | 1936 | -126 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 26 | 1933 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 27 | 1935 | -125.9 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 28 | 1933 | -125.7 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 29 | 1934 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
70 | 20 | NXB | TRUE | 0 | NXB70 | TRUE | TRUE | 30 | 1936 | -125.8 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 1 | 4163 | -273.2 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 2 | 4165 | -273.2 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 3 | 4164 | -273.4 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 4 | 4164 | -273.1 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 5 | 4162 | -273.2 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 6 | 4163 | -273 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 7 | 4162 | -273.4 | 100.04 | 22.545773 | 22.538497 | 23.027758 | 58.341593 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 8 | 4164 | -273.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 9 | 4164 | -273.2 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 10 | 4163 | -273.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 11 | 4164 | -273.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 12 | 4160 | -273.2 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 13 | 4164 | -273.7 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 14 | 4161 | -273.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 15 | 4163 | -273.6 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 16 | 4166 | -273.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 17 | 4163 | -273.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 18 | 4166 | -273.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 19 | 4162 | -273.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 20 | 4163 | -273.2 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 21 | 4163 | -272.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 22 | 4164 | -273.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 23 | 4164 | -273.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 24 | 4163 | -273.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 25 | 4164 | -273.7 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 26 | 4162 | -273 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 27 | 4164 | -273.6 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 28 | 4163 | -272.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 29 | 4164 | -273.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
100 | 20 | NXB | TRUE | 0 | NXB100 | TRUE | TRUE | 30 | 4166 | -273.7 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 1 | 5899 | -393.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 2 | 5902 | -393.6 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 3 | 5899 | -393.1 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 4 | 5900 | -394 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 5 | 5900 | -392.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 6 | 5900 | -393.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 7 | 5900 | -393.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 8 | 5900 | -392.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 9 | 5902 | -392.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 10 | 5899 | -393.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 11 | 5902 | -393.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 12 | 5900 | -393.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 13 | 5898 | -393.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 14 | 5904 | -392.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 15 | 5898 | -393.1 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 16 | 5904 | -393.5 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 17 | 5899 | -393.2 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 18 | 5903 | -393.7 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 19 | 5901 | -393.3 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 20 | 5899 | -393.7 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 21 | 5901 | -393.4 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 22 | 5899 | -393.2 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 23 | 5903 | -393.6 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 24 | 5900 | -393 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 25 | 5900 | -394 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 26 | 5900 | -393 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 27 | 5900 | -393.7 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 28 | 5899 | -392.9 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 29 | 5899 | -394 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
120 | 20 | NXB | TRUE | 0 | NXB120 | TRUE | TRUE | 30 | 5903 | -393.6 | 100.04 | 22.547461 | 22.536346 | 23.032244 | 58.353865 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 1 | 7777 | -528.1 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 2 | 7779 | -528.3 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 3 | 7778 | -526.6 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 4 | 7779 | -527.6 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 5 | 7781 | -527.5 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 6 | 7779 | -528.1 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 7 | 7782 | -528.1 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 8 | 7778 | -526.9 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 9 | 7778 | -527.3 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 10 | 7779 | -526.7 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 11 | 7779 | -527.6 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 12 | 7778 | -527.2 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 13 | 7780 | -528.1 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 14 | 7781 | -527.9 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 15 | 7777 | -527.9 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 16 | 7781 | -528 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 17 | 7778 | -527.4 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 18 | 7780 | -528 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 19 | 7782 | -527.9 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 20 | 7780 | -527.9 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 21 | 7782 | -527.9 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 22 | 7779 | -527.3 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 23 | 7778 | -527.8 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 24 | 7779 | -527.5 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 25 | 7779 | -527.3 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 26 | 7781 | -527.5 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 27 | 7780 | -527.4 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 28 | 7780 | -527.7 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 29 | 7778 | -527.7 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
140 | 20 | NXB | TRUE | 0 | NXB140 | TRUE | TRUE | 30 | 7778 | -527.8 | 100.042 | 22.549656 | 22.535394 | 23.035859 | 58.301789 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 1 | 1462 | -90.15 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 2 | 1460 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 3 | 1461 | -90.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 4 | 1460 | -90.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 5 | 1459 | -90.35 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 6 | 1460 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 7 | 1459 | -90.15 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 8 | 1460 | -89.9 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 9 | 1459 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 10 | 1460 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 11 | 1459 | -90.05 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 12 | 1460 | -90.3 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 13 | 1459 | -90.05 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 14 | 1459 | -90.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 15 | 1461 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 16 | 1459 | -89.95 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 17 | 1461 | -90.35 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 18 | 1459 | -90.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 19 | 1460 | -90.15 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 20 | 1460 | -90.25 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 21 | 1459 | -90 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 22 | 1461 | -90.05 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 23 | 1459 | -90.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 24 | 1461 | -89.9 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 25 | 1459 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 26 | 1459 | -90 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 27 | 1459 | -90.05 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 28 | 1459 | -90.05 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 29 | 1460 | -90 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
70 | 20 | NXC | TRUE | 0 | NXC70 | TRUE | TRUE | 30 | 1459 | -90.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 1 | 3352 | -212.3 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 2 | 3349 | -212 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 3 | 3351 | -212 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 4 | 3350 | -211.8 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 5 | 3351 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 6 | 3351 | -212.4 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 7 | 3350 | -212 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 8 | 3352 | -212.3 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 9 | 3349 | -211.8 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 10 | 3350 | -212 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 11 | 3349 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 12 | 3351 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 13 | 3350 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 14 | 3349 | -211.6 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 15 | 3353 | -212.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 16 | 3350 | -212.4 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 17 | 3351 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 18 | 3350 | -212.2 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 19 | 3351 | -212.3 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 20 | 3351 | -212 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 21 | 3351 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 22 | 3352 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 23 | 3350 | -212.1 | 100.042 | 22.555183 | 22.540529 | 23.040003 | 58.281219 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 24 | 3353 | -212.3 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 25 | 3350 | -212.3 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 26 | 3350 | -212.4 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 27 | 3351 | -212.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 28 | 3350 | -212.1 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 29 | 3351 | -212.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
100 | 20 | NXC | TRUE | 0 | NXC100 | TRUE | TRUE | 30 | 3350 | -212.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 1 | 4870 | -315.4 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 2 | 4868 | -315 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 3 | 4870 | -315.4 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 4 | 4870 | -315 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 5 | 4868 | -315 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 6 | 4871 | -315 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 7 | 4868 | -315.3 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 8 | 4871 | -315.1 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 9 | 4867 | -315.4 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 10 | 4869 | -315.3 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 11 | 4869 | -314.7 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 12 | 4869 | -315.3 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 13 | 4868 | -314.6 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 14 | 4869 | -314.9 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 15 | 4870 | -315.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 16 | 4867 | -315 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 17 | 4869 | -315.3 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 18 | 4866 | -315.1 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 19 | 4868 | -315.5 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 20 | 4872 | -315.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 21 | 4868 | -315.1 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 22 | 4870 | -315.1 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 23 | 4867 | -315.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 24 | 4869 | -314.8 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 25 | 4868 | -314.9 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 26 | 4868 | -314.7 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 27 | 4870 | -314.8 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 28 | 4869 | -315.2 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 29 | 4868 | -315.6 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
120 | 20 | NXC | TRUE | 0 | NXC120 | TRUE | TRUE | 30 | 4867 | -315.4 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 1 | 6537 | -432.6 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 2 | 6535 | -432.8 | 100.042 | 22.560673 | 22.547978 | 23.042885 | 58.299482 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 3 | 6536 | -432.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 4 | 6539 | -432.2 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 5 | 6537 | -432.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 6 | 6539 | -433 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 7 | 6535 | -432.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 8 | 6538 | -433.2 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 9 | 6537 | -432.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 10 | 6537 | -432.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 11 | 6538 | -432.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 12 | 6537 | -432.6 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 13 | 6539 | -432.6 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 14 | 6538 | -432.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 15 | 6539 | -432.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 16 | 6537 | -432.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 17 | 6539 | -433 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 18 | 6539 | -432.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 19 | 6538 | -432.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 20 | 6540 | -432.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 21 | 6536 | -432.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 22 | 6537 | -432.9 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 23 | 6538 | -432.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 24 | 6537 | -432.6 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 25 | 6537 | -432.1 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 26 | 6537 | -432.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 27 | 6539 | -432.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 28 | 6535 | -432.6 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 29 | 6536 | -432.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
140 | 19.9 | NXC | TRUE | 0 | NXC140 | TRUE | TRUE | 30 | 6537 | -432.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 1 | 7427 | -496.6 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 2 | 7430 | -496.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 3 | 7427 | -497 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 4 | 7426 | -496.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 5 | 7427 | -497 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 6 | 7428 | -496.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 7 | 7430 | -496.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 8 | 7427 | -496.9 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 9 | 7431 | -497.1 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 10 | 7427 | -496.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 11 | 7430 | -497.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 12 | 7431 | -497.1 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 13 | 7427 | -497.1 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 14 | 7429 | -497.4 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 15 | 7429 | -496.9 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 16 | 7428 | -497.2 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 17 | 7428 | -496.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 18 | 7430 | -496.9 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 19 | 7427 | -497.3 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 20 | 7428 | -497.1 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 21 | 7430 | -497.5 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 22 | 7427 | -496.9 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 23 | 7433 | -496.6 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 24 | 7430 | -496.7 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 25 | 7432 | -496.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 26 | 7434 | -496.8 | 100.042 | 22.566834 | 22.557737 | 23.046055 | 58.294028 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 27 | 7430 | -497.4 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 28 | 7433 | -497.3 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 29 | 7430 | -497.1 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
150 | 19.9 | NXC | TRUE | 0 | NXC150 | TRUE | TRUE | 30 | 7431 | -497.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 1 | 2350 | -140.6 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 2 | 2348 | -140.4 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 3 | 2350 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 4 | 2348 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 5 | 2348 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 6 | 2349 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 7 | 2350 | -140.8 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 8 | 2351 | -140.9 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 9 | 2349 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 10 | 2350 | -140.8 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 11 | 2350 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 12 | 2349 | -140.8 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 13 | 2349 | -140.6 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 14 | 2348 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 15 | 2350 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 16 | 2347 | -140.9 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 17 | 2349 | -140.6 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 18 | 2348 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 19 | 2348 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 20 | 2350 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 21 | 2349 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 22 | 2350 | -141 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 23 | 2348 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 24 | 2348 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 25 | 2349 | -140.5 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 26 | 2348 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 27 | 2349 | -140.6 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 28 | 2348 | -140.4 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 29 | 2349 | -140.7 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
100 | 20 | NXD | TRUE | 0 | NXD100 | TRUE | TRUE | 30 | 2347 | -140.6 | 100.04 | 22.573683 | 22.569008 | 23.050602 | 58.246064 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 1 | 3553 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 2 | 3552 | -219.7 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 3 | 3553 | -220.1 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 4 | 3555 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 5 | 3552 | -220 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 6 | 3554 | -219.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 7 | 3551 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 8 | 3554 | -220 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 9 | 3553 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 10 | 3553 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 11 | 3553 | -219.5 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 12 | 3553 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 13 | 3553 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 14 | 3551 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 15 | 3553 | -219.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 16 | 3552 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 17 | 3553 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 18 | 3552 | -219.5 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 19 | 3552 | -219.5 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 20 | 3555 | -219 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 21 | 3552 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 22 | 3555 | -219.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 23 | 3552 | -219.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 24 | 3553 | -220 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 25 | 3554 | -220 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 26 | 3553 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 27 | 3555 | -219.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 28 | 3553 | -219.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 29 | 3553 | -219.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
120 | 20 | NXD | TRUE | 0 | NXD120 | TRUE | TRUE | 30 | 3552 | -219.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 1 | 4909 | -312.5 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 2 | 4908 | -312.2 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 3 | 4909 | -312.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 4 | 4911 | -312.1 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 5 | 4909 | -312.2 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 6 | 4911 | -312.8 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 7 | 4909 | -313 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 8 | 4910 | -312.2 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 9 | 4911 | -312.4 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 10 | 4910 | -312.7 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 11 | 4911 | -312.4 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 12 | 4909 | -312.3 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 13 | 4911 | -312.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 14 | 4909 | -312.2 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 15 | 4909 | -312.1 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 16 | 4910 | -312.2 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 17 | 4912 | -312.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 18 | 4914 | -312.6 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 19 | 4912 | -312.5 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 20 | 4913 | -312.9 | 100.038 | 22.574286 | 22.569892 | 23.053686 | 58.219492 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 21 | 4910 | -312.8 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 22 | 4912 | -312.8 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 23 | 4913 | -312.3 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 24 | 4911 | -312.3 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 25 | 4913 | -312.9 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 26 | 4910 | -312.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 27 | 4913 | -312.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 28 | 4911 | -313.1 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 29 | 4912 | -312.3 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
140 | 19.9 | NXD | TRUE | 0 | NXD140 | TRUE | TRUE | 30 | 4912 | -312.7 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 1 | 5648 | -364.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 2 | 5648 | -364.5 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 3 | 5647 | -364.3 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 4 | 5648 | -364.5 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 5 | 5647 | -364.1 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 6 | 5646 | -364.5 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 7 | 5646 | -364.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 8 | 5647 | -364.5 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 9 | 5649 | -364.2 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 10 | 5647 | -364.5 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 11 | 5649 | -364.9 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 12 | 5646 | -364.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 13 | 5649 | -364.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 14 | 5648 | -364.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 15 | 5648 | -364.8 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 16 | 5649 | -364 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 17 | 5647 | -364.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 18 | 5650 | -364.3 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 19 | 5647 | -364.1 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 20 | 5648 | -365 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 21 | 5647 | -364.7 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 22 | 5647 | -364.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 23 | 5647 | -364.2 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 24 | 5649 | -364.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 25 | 5650 | -364.7 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 26 | 5646 | -364.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 27 | 5647 | -364.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 28 | 5646 | -364.4 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 29 | 5648 | -364.6 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
150 | 19.9 | NXD | TRUE | 0 | NXD150 | TRUE | TRUE | 30 | 5647 | -364.5 | 100.037 | 22.579917 | 22.572508 | 23.058009 | 58.219057 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 1 | 7382 | -504 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 2 | 7383 | -504 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 3 | 7383 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 4 | 7385 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 5 | 7383 | -503.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 6 | 7383 | -503.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 7 | 7382 | -503.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 8 | 7383 | -504.5 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 9 | 7381 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 10 | 7381 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 11 | 7384 | -504.5 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 12 | 7383 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 13 | 7385 | -504.8 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 14 | 7380 | -504.3 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 15 | 7382 | -504.2 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 16 | 7384 | -503.9 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 17 | 7382 | -504.2 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 18 | 7385 | -505.1 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 19 | 7382 | -504.1 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 20 | 7383 | -504.7 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 21 | 7383 | -504.3 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 22 | 7383 | -504.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 23 | 7383 | -503.9 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 24 | 7382 | -503.9 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 25 | 7383 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 26 | 7382 | -504.2 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 27 | 7382 | -504.7 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 28 | 7382 | -504.7 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 29 | 7382 | -504.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
200 | 15 | NXD | TRUE | 0 | NXD200 | TRUE | TRUE | 30 | 7382 | -503.8 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 1 | 1952 | -131.7 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 2 | 1953 | -131.9 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 3 | 1951 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 4 | 1951 | -131.4 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 5 | 1953 | -131.2 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 6 | 1953 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 7 | 1954 | -131.5 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 8 | 1952 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 9 | 1952 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 10 | 1953 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 11 | 1952 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 12 | 1952 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 13 | 1952 | -131.6 | 100.038 | 22.58404 | 22.576955 | 23.060836 | 58.203143 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 14 | 1953 | -131.2 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 15 | 1952 | -131.3 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 16 | 1952 | -131.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 17 | 1951 | -131.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 18 | 1953 | -131.5 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 19 | 1954 | -131.3 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 20 | 1952 | -131.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 21 | 1953 | -131.3 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 22 | 1952 | -131.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 23 | 1952 | -131.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 24 | 1951 | -131.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 25 | 1952 | -131.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 26 | 1953 | -131.5 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 27 | 1952 | -131.4 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 28 | 1954 | -131.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 29 | 1952 | -131.5 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
120 | 20 | NXE | TRUE | 0 | NXE120 | TRUE | TRUE | 30 | 1953 | -131.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 1 | 2941 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 2 | 2942 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 3 | 2941 | -205.5 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 4 | 2941 | -205.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 5 | 2943 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 6 | 2941 | -205.2 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 7 | 2942 | -206.4 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 8 | 2941 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 9 | 2941 | -205.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 10 | 2942 | -205.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 11 | 2942 | -205.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 12 | 2941 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 13 | 2942 | -205.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 14 | 2943 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 15 | 2941 | -205.9 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 16 | 2942 | -205.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 17 | 2941 | -205.5 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 18 | 2942 | -205.7 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 19 | 2942 | -206 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 20 | 2941 | -205.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 21 | 2943 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 22 | 2940 | -205.5 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 23 | 2942 | -206.1 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 24 | 2941 | -206 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 25 | 2941 | -205.6 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 26 | 2943 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 27 | 2942 | -205.9 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 28 | 2943 | -205.8 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 29 | 2940 | -206.1 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
140 | 20 | NXE | TRUE | 0 | NXE140 | TRUE | TRUE | 30 | 2943 | -206.2 | 100.04 | 22.590628 | 22.583783 | 23.065334 | 58.158612 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 1 | 3501 | -248.7 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 2 | 3502 | -249.3 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 3 | 3501 | -249 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 4 | 3502 | -248.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 5 | 3504 | -249.1 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 6 | 3502 | -248.8 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 7 | 3504 | -249.3 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 8 | 3502 | -249.4 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 9 | 3502 | -249.2 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 10 | 3503 | -248.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 11 | 3502 | -249.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 12 | 3504 | -248.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 13 | 3501 | -248.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 14 | 3502 | -249.1 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 15 | 3502 | -249.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 16 | 3502 | -249.2 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 17 | 3503 | -248.6 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 18 | 3502 | -248.8 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 19 | 3504 | -249.1 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 20 | 3501 | -249 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 21 | 3503 | -248.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 22 | 3502 | -248.8 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 23 | 3503 | -249.4 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 24 | 3502 | -248.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 25 | 3503 | -249.3 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 26 | 3504 | -248.7 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 27 | 3501 | -249.2 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 28 | 3504 | -249.2 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 29 | 3503 | -248.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
150 | 20 | NXE | TRUE | 0 | NXE150 | TRUE | TRUE | 30 | 3504 | -249.3 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 1 | 5161 | -390.8 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 2 | 5161 | -390.4 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 3 | 5158 | -390.3 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 4 | 5158 | -390.4 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 5 | 5162 | -390.9 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 6 | 5160 | -390.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 7 | 5161 | -390.6 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 8 | 5160 | -390.8 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 9 | 5160 | -391.2 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 10 | 5161 | -390.8 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 11 | 5160 | -390.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 12 | 5161 | -390.4 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 13 | 5161 | -390.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 14 | 5162 | -390.4 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 15 | 5160 | -390.5 | 100.038 | 22.594469 | 22.587526 | 23.06836 | 58.141345 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 16 | 5160 | -390.9 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 17 | 5162 | -390.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 18 | 5160 | -390.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 19 | 5161 | -390.6 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 20 | 5160 | -390 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 21 | 5160 | -390.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 22 | 5159 | -390.2 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 23 | 5160 | -391.2 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 24 | 5159 | -390.8 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 25 | 5159 | -390.8 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 26 | 5162 | -390.3 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 27 | 5159 | -390.9 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 28 | 5161 | -390.9 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 29 | 5159 | -390.8 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
200 | 14.9 | NXE | TRUE | 0 | NXE200 | TRUE | TRUE | 30 | 5161 | -390.3 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 1 | 6715 | -530.2 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 2 | 6715 | -530.1 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 3 | 6715 | -529.7 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 4 | 6714 | -530.6 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 5 | 6716 | -530.4 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 6 | 6715 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 7 | 6715 | -530.7 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 8 | 6714 | -530.3 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 9 | 6716 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 10 | 6716 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 11 | 6714 | -530.1 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 12 | 6716 | -530.8 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 13 | 6715 | -530 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 14 | 6716 | -530.1 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 15 | 6717 | -530.6 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 16 | 6716 | -529.9 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 17 | 6716 | -530.2 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 18 | 6716 | -530.3 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 19 | 6716 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 20 | 6714 | -530.7 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 21 | 6716 | -530.7 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 22 | 6714 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 23 | 6716 | -530.6 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 24 | 6715 | -530.1 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 25 | 6714 | -530.1 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 26 | 6716 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 27 | 6714 | -530 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 28 | 6715 | -530.8 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 29 | 6714 | -530 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
250 | 12 | NXE | TRUE | 0 | NXE250 | TRUE | TRUE | 30 | 6716 | -530.5 | 100.038 | 22.59687 | 22.590429 | 23.07138 | 58.158773 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 1 | 1487 | -107.1 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 2 | 1486 | -106.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 3 | 1487 | -106.6 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 4 | 1488 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 5 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 6 | 1488 | -106.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 7 | 1487 | -106.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 8 | 1487 | -107 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 9 | 1486 | -106.6 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 10 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 11 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 12 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 13 | 1487 | -106.6 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 14 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 15 | 1486 | -106.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 16 | 1486 | -106.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 17 | 1487 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 18 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 19 | 1486 | -107 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 20 | 1488 | -106.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 21 | 1486 | -106.6 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 22 | 1487 | -107 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 23 | 1486 | -106.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 24 | 1487 | -106.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 25 | 1487 | -106.6 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 26 | 1486 | -106.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 27 | 1488 | -106.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 28 | 1487 | -106.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 29 | 1486 | -106.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
140 | 16.8 | NXF | TRUE | 0 | NXF140 | TRUE | TRUE | 30 | 1487 | -106.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 1 | 2170 | -159 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 2 | 2169 | -158.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 3 | 2170 | -158.9 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 4 | 2171 | -158.6 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 5 | 2170 | -158.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 6 | 2170 | -158.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 7 | 2169 | -159 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 8 | 2170 | -159 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 9 | 2170 | -158.8 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 10 | 2170 | -158.7 | 100.039 | 22.602537 | 22.596868 | 23.07608 | 58.16794 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 11 | 2171 | -158.7 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 12 | 2169 | -158.7 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 13 | 2172 | -158.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 14 | 2170 | -159 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 15 | 2170 | -158.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 16 | 2170 | -158.6 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 17 | 2170 | -158.5 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 18 | 2171 | -158.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 19 | 2169 | -158.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 20 | 2171 | -159 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 21 | 2170 | -158.5 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 22 | 2169 | -158.7 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 23 | 2170 | -158.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 24 | 2170 | -158.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 25 | 2172 | -158.6 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 26 | 2170 | -158.7 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 27 | 2171 | -158.9 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 28 | 2169 | -158.7 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 29 | 2170 | -158.6 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
150 | 20 | NXF | TRUE | 0 | NXF150 | TRUE | TRUE | 30 | 2171 | -158.7 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 1 | 3550 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 2 | 3550 | -277.5 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 3 | 3549 | -277.5 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 4 | 3551 | -277.1 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 5 | 3549 | -277 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 6 | 3550 | -277.1 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 7 | 3549 | -277.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 8 | 3550 | -277.5 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 9 | 3551 | -277.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 10 | 3548 | -277.5 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 11 | 3550 | -277.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 12 | 3551 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 13 | 3550 | -276.9 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 14 | 3550 | -276.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 15 | 3550 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 16 | 3551 | -277.4 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 17 | 3550 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 18 | 3550 | -277.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 19 | 3549 | -276.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 20 | 3550 | -277.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 21 | 3550 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 22 | 3549 | -277.3 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 23 | 3552 | -277.4 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 24 | 3549 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 25 | 3551 | -277.2 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 26 | 3549 | -277 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 27 | 3550 | -277.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 28 | 3550 | -277.8 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 29 | 3550 | -277.6 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
200 | 15 | NXF | TRUE | 0 | NXF200 | TRUE | TRUE | 30 | 3551 | -277.4 | 100.04 | 22.606825 | 22.604154 | 23.079188 | 58.210617 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 1 | 4923 | -400.4 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 2 | 4924 | -401.1 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 3 | 4924 | -400.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 4 | 4924 | -400.9 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 5 | 4925 | -401.1 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 6 | 4925 | -400.4 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 7 | 4925 | -400.7 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 8 | 4924 | -400.6 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 9 | 4925 | -401 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 10 | 4924 | -400.9 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 11 | 4925 | -400.7 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 12 | 4925 | -401 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 13 | 4925 | -401.1 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 14 | 4926 | -400.7 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 15 | 4924 | -400.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 16 | 4924 | -401.4 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 17 | 4924 | -400.5 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 18 | 4925 | -401 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 19 | 4926 | -400.6 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 20 | 4924 | -400.2 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 21 | 4926 | -400.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 22 | 4925 | -400.2 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 23 | 4925 | -401.2 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 24 | 4924 | -400.6 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 25 | 4926 | -401.2 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 26 | 4925 | -400.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 27 | 4926 | -400.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 28 | 4924 | -401.4 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 29 | 4925 | -400.6 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
250 | 12 | NXF | TRUE | 0 | NXF250 | TRUE | TRUE | 30 | 4925 | -400.9 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 1 | 5367 | -445.2 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 2 | 5370 | -445.7 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 3 | 5367 | -444.6 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 4 | 5369 | -445.1 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 5 | 5370 | -445.7 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 6 | 5368 | -445.3 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 7 | 5369 | -445.1 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 8 | 5368 | -445.9 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 9 | 5369 | -445.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 10 | 5367 | -445.5 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 11 | 5369 | -445.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 12 | 5368 | -444.8 | 100.039 | 22.613095 | 22.609258 | 23.083764 | 58.179891 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 13 | 5368 | -444.6 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 14 | 5369 | -445.9 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 15 | 5368 | -445.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 16 | 5367 | -444.9 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 17 | 5368 | -444.7 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 18 | 5369 | -445.1 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 19 | 5368 | -445.1 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 20 | 5368 | -445 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 21 | 5370 | -444.6 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 22 | 5368 | -445.5 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 23 | 5368 | -445 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 24 | 5368 | -445.1 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 25 | 5369 | -446 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 26 | 5370 | -445.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 27 | 5369 | -445 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 28 | 5368 | -445.4 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 29 | 5369 | -445.2 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
280 | 10 | NXF | TRUE | 0 | NXF280 | TRUE | TRUE | 30 | 5370 | -445.5 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 1 | 1069 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 2 | 1069 | -79.2 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 3 | 1068 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 4 | 1068 | -79.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 5 | 1069 | -79.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 6 | 1068 | -79.2 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 7 | 1069 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 8 | 1068 | -79.45 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 9 | 1068 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 10 | 1069 | -79.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 11 | 1067 | -79.15 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 12 | 1069 | -79.4 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 13 | 1068 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 14 | 1068 | -79.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 15 | 1069 | -79.1 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 16 | 1068 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 17 | 1068 | -79.5 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 18 | 1068 | -79.2 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 19 | 1069 | -79.2 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 20 | 1068 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 21 | 1069 | -79.5 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 22 | 1068 | -79.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 23 | 1068 | -79.3 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 24 | 1068 | -79.25 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 25 | 1068 | -79.35 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 26 | 1068 | -79.15 | 100.04 | 22.618452 | 22.613181 | 23.086566 | 58.22382 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 27 | 1068 | -79.45 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 28 | 1068 | -79.3 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 29 | 1069 | -79.3 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
150 | 15.1 | NXG | TRUE | 0 | NXG150 | TRUE | TRUE | 30 | 1068 | -79.35 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 1 | 2594 | -206.2 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 2 | 2592 | -205.8 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 3 | 2593 | -205.9 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 4 | 2594 | -206.1 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 5 | 2593 | -206.4 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 6 | 2594 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 7 | 2593 | -205.9 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 8 | 2594 | -205.9 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 9 | 2593 | -205.9 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 10 | 2594 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 11 | 2593 | -205.8 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 12 | 2594 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 13 | 2593 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 14 | 2592 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 15 | 2593 | -205.9 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 16 | 2593 | -206.2 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 17 | 2594 | -206.2 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 18 | 2593 | -206.2 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 19 | 2592 | -206.1 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 20 | 2594 | -206.1 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 21 | 2593 | -205.9 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 22 | 2595 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 23 | 2593 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 24 | 2594 | -206.3 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 25 | 2593 | -205.7 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 26 | 2593 | -205.8 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 27 | 2594 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 28 | 2592 | -206 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 29 | 2594 | -206.1 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
200 | 15 | NXG | TRUE | 0 | NXG200 | TRUE | TRUE | 30 | 2593 | -206.2 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 1 | 3798 | -313.7 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 2 | 3797 | -314 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 3 | 3797 | -313.7 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 4 | 3799 | -313.5 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 5 | 3798 | -313.5 | 100.04 | 22.627847 | 22.619041 | 23.08978 | 58.311987 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 6 | 3800 | -313.7 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 7 | 3797 | -314.2 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 8 | 3799 | -314.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 9 | 3797 | -313.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 10 | 3798 | -313.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 11 | 3798 | -313.8 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 12 | 3798 | -314 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 13 | 3798 | -313.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 14 | 3797 | -313.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 15 | 3799 | -313.7 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 16 | 3798 | -313.5 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 17 | 3798 | -313.8 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 18 | 3799 | -313.4 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 19 | 3798 | -314.3 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 20 | 3799 | -314.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 21 | 3797 | -313.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 22 | 3798 | -313.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 23 | 3798 | -313.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 24 | 3799 | -313.8 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 25 | 3798 | -313.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 26 | 3797 | -314 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 27 | 3799 | -313.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 28 | 3798 | -313.4 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 29 | 3798 | -313.8 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 30 | 3798 | -313.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 1 | 4244 | -357.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 2 | 4245 | -357 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 3 | 4245 | -356.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 4 | 4245 | -356.5 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 5 | 4244 | -356.5 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 6 | 4245 | -357 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 7 | 4244 | -356.7 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 8 | 4245 | -356.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 9 | 4245 | -356.3 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 10 | 4245 | -357.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 11 | 4245 | -357 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 12 | 4245 | -356.7 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 13 | 4245 | -357.2 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 14 | 4245 | -357 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 15 | 4244 | -356.7 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 16 | 4244 | -356.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 17 | 4245 | -356.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 18 | 4246 | -356.7 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 19 | 4245 | -357 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 20 | 4246 | -356.6 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 21 | 4245 | -356.8 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 22 | 4246 | -356.8 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 23 | 4245 | -356.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 24 | 4244 | -357.1 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 25 | 4246 | -356.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 26 | 4244 | -357 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 27 | 4245 | -356.9 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 28 | 4246 | -357.2 | 100.039 | 22.636437 | 22.628688 | 23.09457 | 58.367753 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 29 | 4245 | -357.1 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
280 | 10 | NXG | TRUE | 0 | NXG280 | TRUE | TRUE | 30 | 4245 | -356.8 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 1 | 5048 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 2 | 5049 | -428.4 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 3 | 5049 | -428.4 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 4 | 5050 | -428.8 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 5 | 5047 | -429 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 6 | 5048 | -428.8 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 7 | 5048 | -428.3 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 8 | 5049 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 9 | 5049 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 10 | 5048 | -428.7 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 11 | 5048 | -428.4 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 12 | 5048 | -428.7 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 13 | 5050 | -429 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 14 | 5049 | -428.2 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 15 | 5049 | -428.4 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 16 | 5049 | -428.3 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 17 | 5049 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 18 | 5049 | -429 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 19 | 5048 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 20 | 5049 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 21 | 5048 | -428.6 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 22 | 5047 | -428.4 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 23 | 5047 | -428.4 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 24 | 5048 | -428.5 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 25 | 5049 | -429 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 26 | 5047 | -428.6 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 27 | 5050 | -429.2 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 28 | 5048 | -427.8 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 29 | 5049 | -429 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
300 | 10 | NXG | TRUE | 0 | NXG300 | TRUE | TRUE | 30 | 5050 | -428.7 | 100.038 | 22.639811 | 22.632239 | 23.097586 | 58.178317 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 1 | 1441 | -114.8 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 2 | 1441 | -114.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 3 | 1440 | -114.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 4 | 1441 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 5 | 1441 | -114.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 6 | 1441 | -114.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 7 | 1442 | -114.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 8 | 1441 | -114.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 9 | 1441 | -114.8 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 10 | 1440 | -114.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 11 | 1441 | -114.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 12 | 1441 | -114.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 13 | 1441 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 14 | 1441 | -114.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 15 | 1440 | -114.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 16 | 1441 | -114.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 17 | 1440 | -114.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 18 | 1441 | -114.8 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 19 | 1440 | -114.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 20 | 1441 | -114.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 21 | 1442 | -114.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 22 | 1441 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 23 | 1441 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 24 | 1441 | -114.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 25 | 1442 | -114.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 26 | 1441 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 27 | 1441 | -114.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 28 | 1441 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 29 | 1441 | -114.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
200 | 14 | NXH | TRUE | 0 | NXH200 | TRUE | TRUE | 30 | 1442 | -114.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 1 | 2505 | -206.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 2 | 2507 | -206.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 3 | 2505 | -206 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 4 | 2505 | -206.8 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 5 | 2506 | -206.2 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 6 | 2506 | -206.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 7 | 2506 | -206.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 8 | 2506 | -206.2 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 9 | 2507 | -206.2 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 10 | 2506 | -206.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 11 | 2506 | -206.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 12 | 2505 | -206.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 13 | 2506 | -206.7 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 14 | 2506 | -206.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 15 | 2506 | -206.6 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 16 | 2506 | -206.1 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 17 | 2505 | -206.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 18 | 2506 | -206.5 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 19 | 2505 | -206.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 20 | 2505 | -206.3 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 21 | 2507 | -206.4 | 100.035 | 22.643524 | 22.637417 | 23.102126 | 58.188027 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 22 | 2506 | -206.4 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 23 | 2506 | -206.2 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 24 | 2505 | -206.3 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 25 | 2506 | -206.5 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 26 | 2507 | -206.4 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 27 | 2506 | -206.3 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 28 | 2506 | -206.6 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 29 | 2506 | -206.7 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
250 | 12 | NXH | TRUE | 0 | NXH250 | TRUE | TRUE | 30 | 2506 | -206.7 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 1 | 2918 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 2 | 2919 | -244.3 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 3 | 2919 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 4 | 2919 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 5 | 2920 | -243.8 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 6 | 2920 | -243.9 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 7 | 2921 | -244.1 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 8 | 2919 | -243.9 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 9 | 2920 | -244.2 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 10 | 2920 | -244.6 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 11 | 2920 | -244.1 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 12 | 2920 | -244.1 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 13 | 2920 | -244.2 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 14 | 2920 | -244.4 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 15 | 2919 | -244.2 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 16 | 2919 | -244.4 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 17 | 2920 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 18 | 2920 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 19 | 2920 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 20 | 2920 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 21 | 2920 | -243.6 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 22 | 2919 | -244.1 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 23 | 2920 | -244.2 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 24 | 2919 | -243.8 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 25 | 2919 | -244 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 26 | 2920 | -244.2 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 27 | 2919 | -244.1 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 28 | 2920 | -243.8 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 29 | 2919 | -244.3 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 30 | 2919 | -244.1 | 100.034 | 22.6469 | 22.647368 | 23.105102 | 58.128506 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 1 | 3552 | -299.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 2 | 3552 | -299.2 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 3 | 3552 | -299.2 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 4 | 3552 | -299.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 5 | 3553 | -299.6 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 6 | 3552 | -299.8 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 7 | 3553 | -299.9 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 8 | 3552 | -299.2 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 9 | 3552 | -299.2 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 10 | 3552 | -299.9 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 11 | 3552 | -299.7 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 12 | 3553 | -300.1 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 13 | 3552 | -299.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 14 | 3553 | -299.6 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 15 | 3552 | -298.9 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 16 | 3552 | -299.1 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 17 | 3551 | -299.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 18 | 3552 | -299.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 19 | 3552 | -300.2 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 20 | 3552 | -298.7 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 21 | 3552 | -299.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 22 | 3551 | -299.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 23 | 3551 | -299.6 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 24 | 3552 | -299.1 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 25 | 3552 | -299.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 26 | 3552 | -299.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 27 | 3552 | -299.8 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 28 | 3552 | -299.7 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 29 | 3552 | -299.6 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 30 | 3553 | -299.7 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 1 | 3418 | -290.1 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 2 | 3418 | -290.9 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 3 | 3418 | -290.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 4 | 3417 | -290.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 5 | 3418 | -290.8 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 6 | 3418 | -290.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 7 | 3418 | -290.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 8 | 3417 | -290.8 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 9 | 3417 | -291 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 10 | 3419 | -290.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 11 | 3417 | -290.2 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 12 | 3418 | -290.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 13 | 3417 | -290.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 14 | 3419 | -290.8 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 15 | 3418 | -290.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 16 | 3417 | -290.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 17 | 3418 | -290.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 18 | 3418 | -290.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 19 | 3418 | -290.3 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 20 | 3417 | -291.1 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 21 | 3418 | -290.5 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 22 | 3418 | -290.4 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 23 | 3418 | -290.9 | 100.035 | 22.652903 | 22.654947 | 23.109402 | 58.080745 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 24 | 3417 | -290.6 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 25 | 3417 | -290.3 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 26 | 3417 | -289.7 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 27 | 3417 | -290.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 28 | 3419 | -290.5 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 29 | 3417 | -290.7 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
320 | 8 | NXH | TRUE | 0 | NXH320 | TRUE | TRUE | 30 | 3419 | -290.8 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 1 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 2 | 1646 | -129.9 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 3 | 1646 | -129.8 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 4 | 1646 | -130.2 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 5 | 1646 | -129.9 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 6 | 1645 | -129.9 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 7 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 8 | 1646 | -130 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 9 | 1646 | -130 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 10 | 1646 | -130.3 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 11 | 1646 | -130 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 12 | 1646 | -129.8 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 13 | 1646 | -129.9 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 14 | 1647 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 15 | 1645 | -130.2 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 16 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 17 | 1646 | -129.9 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 18 | 1646 | -129.9 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 19 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 20 | 1645 | -130 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 21 | 1645 | -130 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 22 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 23 | 1646 | -130.3 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 24 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 25 | 1646 | -130.1 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 26 | 1646 | -129.8 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 27 | 1646 | -129.8 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 28 | 1646 | -130.2 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 29 | 1646 | -129.8 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
250 | 12 |  | TRUE | 0 | NXI250 | FALSE | TRUE | 30 | 1646 | -130.4 | 100.036 | 22.658213 | 22.658583 | 23.113096 | 57.994789 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 1 | 1987 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 2 | 1987 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 3 | 1989 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 4 | 1987 | -159.3 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 5 | 1987 | -159.7 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 6 | 1988 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 7 | 1988 | -159.2 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 8 | 1988 | -159.5 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 9 | 1988 | -159.3 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 10 | 1989 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 11 | 1987 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 12 | 1988 | -159.1 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 13 | 1988 | -159.3 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 14 | 1988 | -159.6 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 15 | 1988 | -159.1 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 16 | 1988 | -159.2 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 17 | 1988 | -159.1 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 18 | 1987 | -159.1 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 19 | 1988 | -159.3 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 20 | 1988 | -159.6 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 21 | 1987 | -159.6 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 22 | 1987 | -159.5 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 23 | 1988 | -159.5 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 24 | 1988 | -159.5 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 25 | 1987 | -159.2 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 26 | 1988 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 27 | 1988 | -159.3 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 28 | 1988 | -159.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 29 | 1987 | -159.5 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
280 | 10 |  | TRUE | 0 | NXI280 | FALSE | TRUE | 30 | 1988 | -159.6 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 1 | 2472 | -199.9 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 2 | 2472 | -200 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 3 | 2472 | -199.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 4 | 2472 | -200.4 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 5 | 2472 | -200 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 6 | 2472 | -199.7 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 7 | 2472 | -200.2 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 8 | 2473 | -200 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 9 | 2471 | -199.7 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 10 | 2472 | -199.5 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 11 | 2472 | -199.9 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 12 | 2473 | -199.9 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 13 | 2472 | -200.2 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 14 | 2472 | -199.6 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 15 | 2473 | -200 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 16 | 2472 | -199.8 | 100.036 | 22.66647 | 22.663361 | 23.11766 | 57.965422 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 17 | 2474 | -199.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 18 | 2472 | -199.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 19 | 2472 | -200.2 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 20 | 2471 | -199.9 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 21 | 2472 | -200 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 22 | 2472 | -200 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 23 | 2472 | -199.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 24 | 2473 | -200.1 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 25 | 2472 | -199.9 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 26 | 2472 | -199.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 27 | 2472 | -199.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 28 | 2473 | -200 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 29 | 2472 | -199.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
300 | 10 |  | TRUE | 0 | NXI300 | FALSE | TRUE | 30 | 2472 | -199.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 1 | 2425 | -197.7 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 2 | 2424 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 3 | 2426 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 4 | 2424 | -197.3 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 5 | 2425 | -197.4 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 6 | 2425 | -197.9 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 7 | 2425 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 8 | 2425 | -197.5 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 9 | 2425 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 10 | 2424 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 11 | 2424 | -197.4 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 12 | 2424 | -198 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 13 | 2425 | -198.1 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 14 | 2425 | -198.1 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 15 | 2424 | -197.9 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 16 | 2425 | -197.9 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 17 | 2425 | -197.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 18 | 2425 | -197.5 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 19 | 2425 | -198.3 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 20 | 2425 | -197.5 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 21 | 2425 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 22 | 2426 | -197.7 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 23 | 2425 | -197.4 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 24 | 2426 | -197.7 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 25 | 2425 | -197.8 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 26 | 2426 | -197.9 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 27 | 2425 | -197.7 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 28 | 2426 | -197.6 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 29 | 2425 | -197 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
320 | 8 |  | TRUE | 0 | NXI320 | FALSE | TRUE | 30 | 2425 | -197.4 | 100.037 | 22.670464 | 22.664352 | 23.120883 | 57.960683 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 1 | 3804 | -314 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 2 | 3804 | -314.4 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 3 | 3804 | -314.3 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 4 | 3804 | -314.6 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 5 | 3806 | -314.7 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 6 | 3803 | -314.6 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 7 | 3805 | -313.7 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 8 | 3803 | -313.9 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 9 | 3804 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 10 | 3805 | -313.9 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 11 | 3804 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 12 | 3805 | -314.3 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 13 | 3804 | -314.4 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 14 | 3805 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 15 | 3804 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 16 | 3804 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 17 | 3804 | -314.9 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 18 | 3805 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 19 | 3804 | -313.7 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 20 | 3804 | -314.1 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 21 | 3804 | -314.7 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 22 | 3803 | -314 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 23 | 3805 | -314.1 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 24 | 3803 | -314.6 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 25 | 3804 | -314.1 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 26 | 3805 | -314.3 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 27 | 3804 | -314.5 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 28 | 3806 | -314.1 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 29 | 3804 | -314.6 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
250 | 12 | NXG | TRUE | 0 | NXG250 | TRUE | TRUE | 30 | 3805 | -314.2 | 100.041 | 22.686561 | 22.68224 | 23.136297 | 57.800382 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 1 | 2919 | -243.8 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 2 | 2919 | -243.9 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 3 | 2920 | -244.4 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 4 | 2918 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 5 | 2917 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 6 | 2918 | -243.8 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 7 | 2918 | -243.9 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 8 | 2919 | -243.9 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 9 | 2918 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 10 | 2919 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 11 | 2917 | -243.9 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 12 | 2918 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 13 | 2917 | -244.3 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 14 | 2918 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 15 | 2918 | -244.3 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 16 | 2918 | -244.5 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 17 | 2920 | -243.9 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 18 | 2918 | -243.7 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 19 | 2920 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 20 | 2919 | -244.4 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 21 | 2918 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 22 | 2918 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 23 | 2919 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 24 | 2919 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 25 | 2918 | -244 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 26 | 2919 | -244.2 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 27 | 2919 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 28 | 2918 | -244.3 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 29 | 2920 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
280 | 10 | NXH | TRUE | 0 | NXH280 | TRUE | TRUE | 30 | 2919 | -244.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 1 | 3551 | -299.3 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 2 | 3552 | -299 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 3 | 3552 | -299.4 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 4 | 3552 | -299 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 5 | 3551 | -299.5 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 6 | 3552 | -299.4 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 7 | 3551 | -300.1 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 8 | 3552 | -299.5 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 9 | 3551 | -299.8 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 10 | 3551 | -299.6 | 100.043 | 22.688735 | 22.68149 | 23.140962 | 57.779288 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 11 | 3552 | -299.5 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 12 | 3551 | -299.6 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 13 | 3553 | -299.5 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 14 | 3552 | -299.3 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 15 | 3552 | -299.2 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 16 | 3552 | -299.6 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 17 | 3552 | -299.5 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 18 | 3551 | -299.3 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 19 | 3551 | -299 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 20 | 3552 | -299.8 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 21 | 3551 | -299.3 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 22 | 3553 | -299.3 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 23 | 3551 | -299 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 24 | 3552 | -299.3 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 25 | 3552 | -299.1 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 26 | 3551 | -299.3 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 27 | 3552 | -299.7 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 28 | 3551 | -299.5 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 29 | 3552 | -299.6 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
300 | 10 | NXH | TRUE | 0 | NXH300 | TRUE | TRUE | 30 | 3550 | -299.2 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 1 | 0.05 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 2 | 0.3 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 3 | 0.1 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 4 | 0.15 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 5 | 0.2 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 6 | 0.05 | -0.1 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 7 | 0.2 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 8 | 0.15 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 9 | 0.25 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 10 | 0.1 | -0.1 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 11 | 0.05 | -0.05 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 12 | 0 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 13 | 0.15 | -0.1 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 14 | 0.3 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 15 | 0.2 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 16 | 0.3 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 17 | 0.05 | 0 | 100.039 | 22.692718 | 22.683695 | 23.144065 | 57.775956 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 18 | 0.3 | -0.05 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 19 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 20 | 0.2 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 21 | 0.15 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 22 | 0.2 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 23 | 0.25 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 24 | 0.25 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 25 | 0.05 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 26 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 27 | 0.3 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 28 | 0.05 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 29 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 30 | 0.1 | -0.15 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 31 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 32 | 0.35 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 33 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 34 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 35 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 36 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 37 | 0.2 | -0.2 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 38 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 39 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 40 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 41 | 0.05 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 42 | -0.05 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 43 | 0.25 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 44 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 45 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 46 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 47 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 48 | 0.05 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 49 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 50 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 51 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 52 | 0 | -0.15 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 53 | 0.15 | -0.05 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 54 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 55 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 56 | 0 | -0.05 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 57 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 58 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 59 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 60 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 61 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 62 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 63 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 64 | 0.35 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 65 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 66 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 67 | 0.3 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 68 | 0.15 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 69 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 70 | 0 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 71 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 72 | 0.25 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 73 | 0.05 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 74 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 75 | 0.15 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 76 | 0.1 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 77 | 0.2 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 78 | 0 | -0.05 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 79 | 0 | -0.05 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 80 | 0.35 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 81 | 0 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 82 | 0.2 | -0.05 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 83 | 0.15 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 84 | 0.1 | -0.1 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 85 | 0.2 | 0 | 100.039 | 22.696551 | 22.690109 | 23.149232 | 57.784134 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 86 | 0.1 | 0 | 100.042 | 22.6965 | 22.690524 | 23.151442 | 57.787573 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 87 | 0.2 | 0 | 100.042 | 22.6965 | 22.690524 | 23.151442 | 57.787573 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 88 | 0.2 | 0 | 100.042 | 22.6965 | 22.690524 | 23.151442 | 57.787573 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 89 | 0.2 | 0 | 100.042 | 22.6965 | 22.690524 | 23.151442 | 57.787573 | MEFAC run
40 | 14 | NXB | FALSE | 0 | NXJ40 | FALSE | TRUE | 90 | 0 | 0 | 100.042 | 22.6965 | 22.690524 | 23.151442 | 57.787573 | MEFAC run

## Results
Filter | NK
NXG250 | 47.7474
NXH280 | 47.8
NXH300 | 47.6868
NXA40 | 47.0298
NXA50 | 46.7658
NXA60 | 46.7543
NXA70 | 46.7561
NXA80 | 46.7847
NXA90 | 46.8438
NXB100 | 46.9257
NXB120 | 47.0054
NXB140 | 47.0342
NXB50 | 46.6562
NXB70 | 46.7242
NXC100 | 47.0049
NXC120 | 47.0567
NXC140 | 47.1335
NXC150 | 47.1545
NXC70 | 46.6909
NXD100 | 47.102
NXD120 | 47.229
NXD140 | 47.2702
NXD150 | 47.2999
NXD200 | 47.3791
NXE120 | 47.5595
NXE140 | 47.5646
NXE150 | 47.5683
NXE200 | 47.6029
NXE250 | 47.6294
NXF140 | 47.7196
NXF150 | 47.6965
NXF200 | 47.7062
NXF250 | 47.7097
NXF280 | 47.7108
NXG150 | 47.7442
NXG200 | 47.7624
NXG280 | 47.7153
NXG300 | 47.663
NXH200 | 47.772
NXH250 | 47.8306
NXH320 | 47.6014
NXI250 | 47.8166
NXI280 | 47.819
NXI300 | 47.7361
NXI320 | 47.6209
NXJ100 | 48.6338
NXJ40 | 51.0802
NXJ50 | 50.2919
NXJ60 | 49.7798
NXJ70 | 49.3807
NXJ80 | 49.0641
NXJ90 | 48.7932
NXK100 | 47.5562
NXK40 | 48.8556
NXK50 | 48.3459
NXK60 | 48.0671
NXK70 | 47.8568
NXK80 | 47.742
NXK90 | 47.6226
NXG250* | 47.6707
NXH280* | 47.8605
NXH300* | 47.8206