## MEXReport
Primary Standards Dosimetry Laboratory, Medical Radiation Services
619 Lower Plenty Road, Yallambie, Victoria 3085, Australia
Tel: +613 9433 2211 Fax: +613 9432 1835
E-mail: psdl@arpansa.gov.au Web: www.arpansa.gov.au
In reply please quote:
None
CAL00001
CALIBRATION REPORT
on a therapy ionisation chamber for
MEDIUM-ENERGY KILOVOLTAGE X-RAYS
Client |  |  |  | ClientA_Name
200 Street Name
Suburb NSW 2020
Ionisation chamber |  |  |  | PTW 30013 5122
Period of tests |  |  |  | 12 Feb 2021 to 12 Feb 2021
Previous calibration |  |  |  | Not previously calibrated at ARPANSA | [11 empty] | IRC-FORM-2100J (ARPANSA-FORM-1856)
Test By |  |  |  | Duncan Butler
Report By |  |  |  | Duncan Butler
Report date |  |  |  | 05 Oct 2021
Direct inquiries to |  |  |  | Chris Oliver
Tel: (03) 9433 2333 |  |  |  | Email: psdl@arpansa.gov.au
Signed: |  |  |  |  | (Authorised Signatory) | [5 empty] | Date:
Chris Oliver, Senior Radiation Scientist, Primary Standards Dosimetry Laboratory
per C-M Larsson, CEO of ARPANSA
This certificate is consistent with the capabilities that are
included in Appendix C of the MRA drawn up by the
|  |  | Accredited for compliance with ISO/IEC 17025 - Calibration. | [7 empty] | CIPM. Under the MRA, all participating institutes
|  |  | The results of the tests, calibrations and/or | [7 empty] | recognize the validity of each other’s calibration and
|  |  | measurements included in this document are | [7 empty] | measurement certificates for the quantities, ranges and
|  |  | traceable to Australian/national standards | [7 empty] | measurement uncertainties specified in Appendix C (for
|  |  | NATA Accredited Laboratory Number: 14433 | [7 empty] | details see www.bipm.org).
Calibration No: CAL00001 | [5 empty] | Calibration No: CAL00001
Calibration No: CAL00001 | [5 empty] | Calibration No: CAL00001 |  |  |  |  | Checked: |  |  |  |  | Page 1 of 6
GENERAL COMMENTS
Preliminary Inspection
– | The ionisation chamber had no obvious damage or faults on receipt.
Calibration Coefficient
– | The calibration coefficient is the number by which the charge from the chamber, in nC, must be multiplied to obtain the air kerma [1]. The calibration factor for the electrometer must also be taken into account when measuring the charge from the chamber.
Calibration Coefficients for Medium-Energy X-ray (MEX) Qualities in Air
– | The calibration coefficients for the chamber for each X-ray beam quality from the Gulmay Comet X-ray generator were determined by comparison with the ARPANSA Medium Energy Free-Air Chamber, which is the Australian primary standard of air kerma for medium energy X-rays.
– | The Gulmay Comet X-ray generator is constant potential and the X-ray tube has a tungsten target.
Recombination Correction Measurement
– | Not measured.
Polarity Correction Measurement
– | Not measured.
Notes
– | The ionisation chamber was tested in accordance with ARPANSA Standard Operational Procedure ARPANSA-SOP-0816 Version 7.
References
[1] | AAPM protocol for 40–300 kV x-ray beam dosimetry in radiotherapy and radiobiology, C.-M. Ma, Chair, C. W. Coffey, L. A. DeWerd, C. Liu, R. Nath, S. M. Seltzer, J. P. Seuntjens, Med. Phys. 28 (6) 868-893, 2001
Calibration No: CAL00001 | [5 empty] | Calibration No: CAL00001 |  |  |  |  | Checked: |  |  |  |  | Page 2 of 6
ClientA_Name
PTW 30013 5122
-250
Air Kerma Calibration Certificate - Medium-Energy X-rays
Client |  |  |  | ClientA_Name
Ionisation chamber |  |  |  | PTW 30013 5122
Polarising voltage |  |  |  | -250
Collected charge polarity |  |  |  | (-250) on the guard electrode Positive (Central Electrode Negative)
Reference point |  |  |  | The geometrical centre of the cavity
Geometry |  |  |  | Mark on chamber stem facing the radiation source
Chamber stem vertically upwards, cable down
Horizontal radiation beam
12 Feb 2021 to 12 Feb 2021
Circular beam of diameter 10 cm
Build-up |  |  |  | Build-up cap removed except where stated. Calibrated free in air.
Polarity and recombination |  |  |  | Corrections not applied
Reference conditions |  |  |  | 20oC, 101.325 kPa and 50% humidity
Measurement date(s) |  |  |  | 12 Feb 2021 to 12 Feb 2021
NXA50 |  | 50 | 4 |  |  | 2.39 |  | 0.08 |  | 29.93 |  | 0.4 |  | 46.63 | 1.4
NXA70 |  | 70 | 4 |  |  | 3.19 |  | 0.11 |  | 34.27 |  | 0.8 |  | 46.69 | 1.4
NXB100 |  | 100 | 4.5 |  |  | 4.74 | Table 1: Subset of air kerma calibration coefficients | 0.18 |  | 41.59 |  | 1.6 |  | 46.89 | 1.4
NXC120 |  | 120 | 6 |  |  | 6.38 |  | 0.28 |  | 48.95 |  | 1.8 |  | 47.03 | 1.4
NXD140 |  | 140 | 9 |  |  | 8.44 |  | 0.45 |  | 58.06 |  | 1.8 |  | 47.24 | 1.4
NXE150 |  | 150 | 4 | 0.5 |  |  |  | 0.84 |  | 72.15 |  | 1.4 |  | 47.53 | 1.4
NXF200 |  | 200 | 4 | 1 |  |  |  | 1.63 |  | 94.74 |  | 1.6 |  | 47.67 | 1.4
NXG250 |  | 250 | 4 | 1.6 |  |  |  | 2.57 |  | 120.1 |  | 1.8 |  | 47.72 | 1.4
NXH280 |  | 280 | 4 | 3 |  |  |  | 3.5 |  | 147.05 |  | 1.4 |  | 47.76 | 1.4
NXH300 |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.66 | 1.4
NXH300* |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.79 | 1.4
NXF200 |  | 200 | 4 | 1 |  |  |  | 1.63 |  | 94.74 |  | 1.6 |  | 47.67 | 1.4
NXG250 |  | 250 | 4 | 1.6 |  |  |  | 2.57 |  | 120.1 |  | 1.8 |  | 47.72 | 1.4
NXH280 |  | 280 | 4 | 3 |  |  |  | 3.5 |  | 147.05 |  | 1.4 |  | 47.76 | 1.4
NXH300 |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.66 | 1.4
NXH300* |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.79 | 1.4
[1] The energy of a monoenergetic beam with the same HVL in mm of Cu
[2] The air kerma calibration coefficient
* With buildup cap on
Calibrated by: ClientA_Name
Calibration No: CAL00001 | [5 empty] | Calibration No: CAL00001 |  |  |  |  | Checked: |  |  |  |  | Page 3 of 6
NXJ40 |  | 40 | 0.5 |  |  | 0.57 |  |  |  | 17.96 |  | 1.8 |  | 51.05 | 1.4
NXK40 |  | 40 | 1 |  |  | 0.93 | Table 2: Complete set of air kerma calibration coefficients for all MEX beams |  |  | 20.93 |  | 1.1 |  | 48.8 | 1.4
NXA40 |  | 40 | 4 |  |  | 1.9 |  | 0.06 |  | 27.07 |  | 0.2 |  | 46.76 | 1.4
NXJ50 |  | 50 | 0.5 |  |  | 0.68 |  |  |  | 19.5 |  | 2.7 |  | 50.27 | 1.4
NXK50 |  | 50 | 1 |  |  | 1.12 |  |  |  | 22.65 |  | 1.6 |  | 48.31 | 1.4
NXA50 |  | 50 | 4 |  |  | 2.39 |  | 0.08 |  | 29.93 |  | 0.4 |  | 46.63 | 1.4
NXB50 |  | 50 | 4.5 |  |  | 2.53 |  | 0.08 |  | 30.62 |  | 0.3 |  | 46.5 | 1.4
NXJ60 |  | 60 | 0.5 |  |  | 0.79 |  |  |  | 20.86 |  | 3.5 |  | 49.76 | 1.4
NXK60 |  | 60 | 1 |  |  | 1.29 |  |  |  | 24.09 |  | 2.1 |  | 48.04 | 1.4
NXA60 |  | 60 | 4 |  |  | 2.81 |  | 0.1 |  | 32.25 |  | 0.6 |  | 46.67 | 1.4
NXJ70 |  | 70 | 0.5 |  |  | 0.9 |  |  |  | 22.15 |  | 4.1 |  | 49.37 | 1.4
NXK70 |  | 70 | 1 |  |  | 1.46 |  |  |  | 25.44 |  | 2.6 |  | 47.84 | 1.4
NXA70 |  | 70 | 4 |  |  | 3.19 |  | 0.11 |  | 34.27 |  | 0.8 |  | 46.69 | 1.4
NXB70 |  | 70 | 4.5 |  |  | 3.39 |  | 0.12 |  | 35.19 |  | 0.7 |  | 46.65 | 1.4
NXC70 |  | 70 | 6 |  |  | 3.95 |  | 0.14 |  | 37.51 |  | 0.5 |  | 46.59 | 1.4
NXJ80 |  | 80 | 0.5 |  |  | 1.02 |  |  |  | 23.48 |  | 4.8 |  | 49.05 | 1.4
NXK80 |  | 80 | 1 |  |  | 1.64 |  |  |  | 26.84 |  | 3.2 |  | 47.73 | 1.4
NXA80 |  | 80 | 4 |  |  | 3.62 |  | 0.13 |  | 36.33 |  | 1.1 |  | 46.74 | 1.4
NXJ90 |  | 90 | 0.5 |  |  | 1.15 |  |  |  | 24.88 |  | 5.5 |  | 48.78 | 1.4
NXK90 |  | 90 | 1 |  |  | 1.84 |  |  |  | 28.32 |  | 3.7 |  | 47.61 | 1.4
NXA90 |  | 90 | 4 |  |  | 4.04 |  | 0.15 |  | 38.38 |  | 1.4 |  | 46.81 | 1.4
NXJ100 |  | 100 | 0.5 |  |  | 1.3 |  |  |  | 26.34 |  | 6.1 |  | 48.63 | 1.4
NXK100 |  | 100 | 1 |  |  | 2.05 |  |  |  | 29.87 |  | 4.3 |  | 47.54 | 1.4
NXB100 |  | 100 | 4.5 |  |  | 4.74 |  | 0.18 |  | 41.59 |  | 1.6 |  | 46.89 | 1.4
NXC100 |  | 100 | 6 |  |  | 5.49 |  | 0.22 |  | 44.58 |  | 1.2 |  | 46.96 | 1.4
NXD100 |  | 100 | 9 |  |  | 6.61 |  | 0.29 |  | 49.08 |  | 0.8 |  | 47.04 | 1.4
NXB120 |  | 120 | 4.5 |  |  | 5.56 |  | 0.23 |  | 45.74 |  | 2.2 |  | 46.98 | 1.4
NXC120 |  | 120 | 6 |  |  | 6.38 |  | 0.28 |  | 48.95 |  | 1.8 |  | 47.03 | 1.4
NXD120 |  | 120 | 9 |  |  | 7.59 |  | 0.37 |  | 53.75 |  | 1.2 |  | 47.19 | 1.4
NXE120 |  | 120 | 4 | 0.5 |  | 10.31 |  | 0.63 |  | 65 |  | 0.7 |  | 47.49 | 1.4
NXB140 |  | 140 | 4.5 |  |  | 6.33 |  | 0.28 |  | 49.87 |  | 3 |  | 47.02 | 1.4
NXC140 |  | 140 | 6 |  |  | 7.2 |  | 0.34 |  | 53.14 |  | 2.5 |  | 47.11 | 1.4
NXD140 |  | 140 | 9 |  |  | 8.44 |  | 0.45 |  | 58.06 |  | 1.8 |  | 47.24 | 1.4
NXE140 |  | 140 | 4 | 0.5 |  |  |  | 0.77 |  | 69.85 |  | 1.2 |  | 47.52 | 1.4
NXF140 |  | 140 | 4 | 1 |  |  |  | 1.03 |  | 78.52 |  | 0.6 |  | 47.64 | 1.4
NXC150 |  | 150 | 6 |  |  | 7.58 |  | 0.38 |  | 55.19 |  | 2.8 |  | 47.14 | 1.4
NXD150 |  | 150 | 9 |  |  | 8.83 |  | 0.49 |  | 60.13 |  | 2.1 |  | 47.28 | 1.4
NXE150 |  | 150 | 4 | 0.5 |  |  |  | 0.84 |  | 72.15 |  | 1.4 |  | 47.53 | 1.4
NXF150 |  | 150 | 4 | 1 |  |  |  | 1.13 |  | 81.36 |  | 0.9 |  | 47.64 | 1.4
NXG150 |  | 150 | 4 | 1.6 |  |  |  | 1.38 |  | 89.07 |  | 0.4 |  | 47.63 | 1.4
NXD150 |  | 150 | 9 |  |  | 8.83 |  | 0.49 |  | 60.13 |  | 2.1 |  | 47.28 | 1.4
NXE150 |  | 150 | 4 | 0.5 |  |  |  | 0.84 |  | 72.15 |  | 1.4 |  | 47.53 | 1.4
NXF150 |  | 150 | 4 | 1 |  |  |  | 1.13 |  | 81.36 |  | 0.9 |  | 47.64 | 1.4
NXG150 |  | 150 | 4 | 1.6 |  |  |  | 1.38 |  | 89.07 |  | 0.4 |  | 47.63 | 1.4
…continued p5
NXD200 |  | 200 | 9 |  |  | 10.53 |  | 0.73 |  | 69.76 | Checked: | 2.9 |  | 47.36 | 1.4 | Page 4 of 6
NXE200 |  | 200 | 4 | 0.5 |  |  |  | 1.21 |  | 82.81 |  | 2.2 |  | 47.58 | 1.4
NXF200 |  | 200 | 4 | 1 |  |  |  | 1.63 |  | 94.74 |  | 1.6 |  | 47.67 | 1.4
NXG200 |  | 200 | 4 | 1.6 |  |  |  | 2 |  | 105.25 |  | 1.2 |  | 47.72 | 1.4
NXH200 |  | 200 | 4 | 3 |  |  |  | 2.54 |  | 120.15 |  | 0.6 |  | 47.69 | 1.4
NXE250 |  | 250 | 4 | 0.5 |  |  |  | 1.61 |  | 92.84 |  | 3 |  | 47.61 | 1.4
NXF250 |  | 250 | 4 | 1 |  |  |  | 2.14 |  | 107.32 |  | 2.3 |  | 47.69 | 1.4
NXG250 |  | 250 | 4 | 1.6 |  |  |  | 2.57 |  | 120.1 |  | 1.8 |  | 47.72 | 1.4
NXH250 |  | 250 | 4 | 3 |  |  |  | 3.16 |  | 137.46 |  | 1.2 |  | 47.79 | 1.4
NXI250 |  | 250 | 4 | 5 |  |  |  | 3.6 |  | 149.51 |  | 0.7 |  | 47.75 | 1.4
NXF280 |  | 280 | 4 | 1 |  |  |  | 2.43 |  | 114.61 |  | 2.5 |  | 47.69 | 1.4
NXG280 |  | 280 | 4 | 1.6 |  |  |  | 2.88 |  | 128.5 |  | 2 |  | 47.69 | 1.4
NXH280 |  | 280 | 4 | 3 |  |  |  | 3.5 |  | 147.05 |  | 1.4 |  | 47.76 | 1.4
NXI280 |  | 280 | 4 | 5 |  |  |  | 3.93 |  | 159.92 |  | 0.9 |  | 47.76 | 1.4
NXG300 |  | 300 | 4 | 1.6 |  |  |  | 3.1 |  | 133.92 |  | 2.4 |  | 47.64 | 1.4
NXH300 |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.66 | 1.4
NXI300 |  | 300 | 4 | 5 |  |  |  | 4.15 |  | 166.58 |  | 1.1 |  | 47.69 | 1.4
NXH320 |  | 320 | 4 | 3 |  |  |  | 3.9 |  | 159.22 |  | 1.6 |  | 47.57 | 1.4
NXI320 |  | 320 | 4 | 5 |  |  |  | 4.34 |  | 173.06 |  | 1.1 |  | 47.58 | 1.4
NXG250* |  | 250 | 4 | 1.6 |  |  |  | 2.57 |  | 120.1 |  | 1.8 |  | 47.64 | 1.4
NXH280* |  | 280 | 4 | 3 |  |  |  | 3.5 |  | 147.05 |  | 1.4 |  | 47.82 | 1.4
NXH300* |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.79 | 1.4
NXI320 |  | 320 | 4 | 5 |  |  |  | 4.34 |  | 173.06 |  | 1.1 |  | 47.58 | 1.4
NXG250* |  | 250 | 4 | 1.6 |  |  |  | 2.57 |  | 120.1 |  | 1.8 |  | 47.64 | 1.4
NXH280* |  | 280 | 4 | 3 |  |  |  | 3.5 |  | 147.05 |  | 1.4 |  | 47.82 | 1.4
NXH300* |  | 300 | 4 | 3 |  |  |  | 3.7 |  | 153.21 |  | 1.7 |  | 47.79 | 1.4
[1] The energy of a monoenergetic beam with the same HVL in mm of Cu
[2] The air kerma calibration coefficient
* With buildup cap on
#REF!
Calibration No: CAL00001 | [5 empty] | Calibration No: CAL00001 |  |  |  |  | Checked: |  |  |  |  | Page 5 of 6
#REF!
#REF!
Calibration No: CAL00001 | [5 empty] | Calibration No: CAL00001 |  |  |  |  | Checked: |  |  |  |  | Page 6 of 6